## TOC
functional
AddROI
AnalysisProgressIndicator
Anonymise DICOM Images
AuthorisationAccess
AuthorisationAccess
BypassLogin
CommandLineInterface
ControlButtonTooltips
CUstomerWorkflowEU
CustomerWorkflowUK
DeleteROI
DesktopLicensing
DetectionLogging
DICOMImageFormat
DisplayImage
FeatureTour
FolderWatcher
FullScreen
GenerateAndViewFeedback
ImageExport
ImageID
ImageMagnifer
ImageSubmit
LegalInformationPage
ManageAllowedIPRanges
ManageUserLogins
ProgressDisplay
RepositionROI
ROIIndicator
SeeWhatsNew
ShellIntegration
TiffImageFormat
ToggleAnnotation
ToothIndicator
UpdateClient
nonfunctional
CariesDetectionPerformance
EncryptTraffic
UserData
UserData
user
AddCariesROI
AllowIPRange
BypassLogin
CommandLineInterface
ControButtonToolTips
CustomerWorkflowEU
CustomerWorkflowUK
DeleteCariesROI
DesktopLicensing
DesktopUsageMetrics
DisplayCariesDetection
DisplayCariesDetection
DisplayImage
EnterFeedback
ExpandImage
FeatureTour
FolderWatcher
ImageExport
ImageID
LegalLabelling
LogDetections
MagnifyImage
ManageUserLogins
ProtectPatientData
RepositionCariesROI
SeeWhatsNew
ShellIntegration
ShowProgress
SubmitImages
ToggleAnnotationVisibility
UpgradeClient
UsernameLogin
ViewAnalysisProgress
ViewFeedback

## ADDROI
AddROI
| Tags: | @FR, @FR008, @UR008, @Analysis
AddCariesROICursorAppearanceChange
| Tags: | @FR008-1, @UR008-1
|  | Given | I have submitted an image
|  | When | click the Add ROI button
|  | Then | the Add ROI button changes appearance
|  | And | the Cursor changes appearance to a crosshair
AddCariesROICursorAppearanceReturn
| Tags: | @FR008-2, @UR008-1
|  | Given | I have selected the Add ROI Button
|  | When | I hit select the Add ROI button again or any other button
|  | Then | the Add ROI button changes appearance back to original
|  | And | the Cursor changes appearance back to default
AddCariesROI
| Tags: | @FR008-3, @UR008-1
|  | Given | I have selected the Add ROI Button
|  | When | I move the cursor to the image and click
|  | Then | a new ROI is display with the arrow point at the mouse click
|  | And | pointing away from the nearest tooth centre
|  | And | the Add ROI button changes appearance back to original
|  | And | the Cursor changes appearance back to default

## ANALYSISPROGRESSINDICATOR
AnalysisProgressIndicator
| Tags: | @FR, @FR024, @UR022, @Analysis, @WebOnly
ShowFileUploadProgress
| Tags: | @FR024-1, @UR022-1, @WebOnly
|  | Given | I have logged in with analysis permissions
|  | When | I submit an image for analysis
|  | Then | I can see a progress bar
|  | And | the progress bar fills from left to right until completely full
|  | When | the progress bar is full
|  | Then | the image is displayed
|  | And | the progress bar is hidden
ShowAnalysisProgress
| Tags: | @FR024-2, @UR022-2
|  | Given | I have submitted an image for analysis
|  | When | the image is displayed
|  | Then | the status of the analysis is displayed over the image
|  | When | the analysis progresses
|  | Then | the status is updated to reflect the stage of the analysis

## ANONYMISEDICOMIMAGES
Anonymise DICOM Images
| Tags: | @FR, @FR011, @UR011, @Security, @WebOnly
AnonymiseDICOMImage
| Tags: | @FR011-1, @UR011-1
|  | When | I submit a DICOM Image
|  | Then | the trace information in the web browser indicates an empty patient name and default birth date
|  | And | the trace information displayed on the web server indicates an empty patient name and default birth date

## AUTHORISATIONACCESS
AuthorisationAccess
| Tags: | @FR, @FR020, @UR017, @Security, @WebOnly
AnonymousCanAccessHomepage
| Tags: | @FR020-1, @UR017-1
|  | Given | I am logged out
|  | And | no Allowed IP Range exists which includes my IP address
|  | When | I access <base url>
|  | Then | the assistdent.net home page is displayed
AnonymousCanAccessTour
| Tags: | @FR020-2, @UR017-1
|  | Given | I am logged out
|  | And | no Allowed IP Range exists which includes my IP address
|  | When | I access <base url>/tour
|  | Then | the assistdent.net tour page is displayed
AnonymousCanAccessLegal
| Tags: | @FR020-3, @UR017-1
|  | Given | I am logged out
|  | And | no Allowed IP Range exists which includes my IP address
|  | When | I access <base url>/legal
|  | Then | the assistdent.net legal page is displayed
AnonymousCanAccessContact
| Tags: | @FR020-4, @UR017-1
|  | Given | I am logged out
|  | And | no Allowed IP Range exists which includes my IP address
|  | When | I access <base url>/contact
|  | Then | the assistdent.net contact page is displayed
AnonymouseCannotAccessAnalysis
| Tags: | @FR020-5, @UR017-1
|  | Given | I am logged out
|  | And | no Allowed IP Range exists which includes my IP address
|  | When | I access <base url>/analysis
|  | Then | the assistdent.net login page is displayed
AnonymouseCannotAccessFeedback
| Tags: | @FR020-6, @UR017-1
|  | Given | I am logged out
|  | And | no Allowed IP Range exists which includes my IP address
|  | When | I access <base url>/feedback/new
|  | Then | the assistdent.net login page is displayed
AnonymouseCannotAccessManageUsers
| Tags: | @FR020-7, @UR017-1
|  | Given | I am logged out
|  | And | no Allowed IP Range exists which includes my IP address
|  | When | I access <base url>/user
|  | Then | the assistdent.net login page is displayed
AnonymousCannotAccessAllowedIPs
| Tags: | @FR020-8, @UR017-1
|  | Given | I am logged out
|  | And | no Allowed IP Range exists which includes my IP address
|  | When | I access <base url>/allowedips
|  | Then | the assistdent.net login page is displayed
UserCanAccessAnalysis
| Tags: | @FR020-9, @UR017-1
|  | Given | I am logged in with a username and password
|  | And | I have analysis permissions
|  | When | I access <base url>/analysis
|  | Then | the assistdent.net analysis page is displayed
UserCanAccessFeedback
| Tags: | @FR020-10, @UR017-1
|  | Given | I am logged in with a username and password
|  | And | I do not have admin permissions
|  | When | I access <base url>/feedback/new
|  | Then | the assistdent.net feedback page is displayed
UserCannotAccessManageUsers
| Tags: | @FR020-11, @UR017-1
|  | Given | I am logged in with a username and password
|  | And | I do not have admin permissions
|  | When | I access <base url>/user
|  | Then | the assistdent.net login page is displayed
UserCannotAccessAllowedIPs
| Tags: | @FR020-12, @UR017-1
|  | Given | I am logged in with a username and password
|  | And | I do not have admin permissions
|  | When | I access <base url>/allowedips
|  | Then | the assistdent.net login page is displayed
AdminCanAccessManageUsers
| Tags: | @FR020-13, @UR017-1
|  | Given | I am logged in with admin permissions
|  | When | I access <base url>/user
|  | Then | the assistdent.net manage users page is displayed
AdminCanAccessAllowedIPs
| Tags: | @FR020-14, @UR017-1
|  | Given | I am logged in with admin permissions
|  | When | I access <base url>/allowedips
|  | Then | the assistdent.net allowed ips page is displayed
UserRefusedAccessAnalysis
| Tags: | @FR020-15, @UR017-2
|  | Given | I am logged in with a username and password
|  | And | I do not have analysis permissions
|  | When | I access <base url>/analysis
|  | Then | the assistdent.net analysis page is not displayed

## AUTHORISATIONACC(1)
AuthorisationAccess
| Tags: | @FR, @FR023, @UR017, @Security, @WebOnly
UserCanAccessLogin
| Tags: | @FR023-1, @UR017-1
|  | Given | I am logged out
|  | When | I click the Login link
|  | Then | I am redirected to the login page
UserCannotLoginWithIncorrectPassword
| Tags: | @FR023-2, @UR017-1
|  | Given | I am logged out
|  | And | I have navigated to the login page
|  | When | I enter my username
|  | And | I enter an invalid password
|  | Then | I am not granted access
|  | And | I remain on the login page
UserCannotLoginWithIncorrectUsername
| Tags: | @FR023-3, @UR017-1
|  | Given | I am logged out
|  | And | I have navigated to the login page
|  | When | I enter an invalid username
|  | And | I enter a valid password for another user
|  | Then | I am not granted access
|  | And | I remain on the login page
UserCanLoginWithCorrectCredentials
| Tags: | @FR023-4, @UR017-1
|  | Given | I am logged out
|  | And | I have navigated to the login page
|  | When | I enter an valid username
|  | And | I enter a valid password
|  | Then | I am redirected to another page
|  | When | I click the Analysis link
|  | Then | I have access to the Analysis page

## BYPASSLOGIN
BypassLogin
| Tags: | @FR, @FR013, @UR013, @Security, @WebOnly
CanAnalyzeFromAllowedIPAddress
| Tags: | @FR013-1, @UR013-1
|  | Given | the current computer IP is specified exclusively by a single Allowed IP Range using exact IP address
|  | And | I am logged out of the application
|  | When | I access the URL <base>/Analysis
|  | Then | the analysis page is displayed
|  | When | I submit an image
|  | Then | the image and results are displayed
CanAnalyzeFromAllowedIPRange
| Tags: | @FR013-2, @UR013-1
|  | Given | the current computer IP is specified exclusively by a single Allowed IP Range within a range of addresses
|  | And | I am logged out of the application
|  | When | I access the URL <base>/Analysis
|  | Then | the analysis page is displayed
|  | When | I submit an image
|  | Then | the image and results are displayed
CantAnalyzeFeomNoAllowedIP
| Tags: | @FR013-3, @UR013-2
|  | Given | the current computer IP is not covered by any Allowed IP Range entry
|  | And | I am logged out of the application
|  | When | I access the URL <base>/Analysis
|  | Then | the login page is displayed
CantAnalyzeFromExpiredAllowedIP
| Tags: | @FR013-4, @UR013-3
|  | Given | the current computer IP is specified exclusively by an Allowed IP Range which is expired
|  | And | I am logged out of the application
|  | When | I access the URL <base>/Analysis
|  | Then | the login page is displayed

## COMMANDLINEINTERFACE
CommandLineInterface
| Tags: | @FR, @FR026, @UR024, @Integration, @DesktopOnly
PassImage
| Tags: | @FR026-1, @UR024-1
|  | When | I request AssistDent from the command line
|  | And | the --image option is set to "image8.tif"
|  | Then | AssistDent starts
|  | And | the image analyzed is "image8.tif"
PassTitle
| Tags: | @FR026-2, @UR024-1
|  | When | I request AssistDent from the command line
|  | And | the --title option is set to "Developer/Dave"
|  | Then | AssistDent starts
|  | And | the window title contains the text "Developer/Dave"

## CONTROLBUTTONTOOLTIPS
ControlButtonTooltips
| Tags: | @FR, @FR010, @UR010, @Labelling, @WebOnly
ShowTooltip
| Tags: | @FR010-1, @UR010-1
|  | Given | an analyzed bitewing image
|  | When | I hover over the <button name> button
|  | Then | a tooltip with the text <tooltip text> appears above the button
Examples
|  |  | button name | tooltip text
|  |  | toggle | Hide
|  |  | delete | Delete
|  |  | flip | Flip
|  |  | add | Add
|  |  | full | Full Screen

## CUSTOMERWORKFLOWEU
CUstomerWorkflowEU
| Tags: | @FR, @FR031, @UR029, @Admin, @DesktopOnly
CanRegisterInCECountry
| Tags: | @FR031-1, @UR029-1
|  | Given | no licence exists for the current machine
|  | And | I start Assistdent
|  | And | the use registration form is displayed
|  | When | I enter valid details
|  | And | I select "Germany" for the country
|  | Then | the registration is successful
|  | And | AssistDent starts
TakePaymentEU:
| Tags: | @FR031-2, @UR029-2
|  | Given | I have register a trial licence for an EU user
|  | When | I access the purchase licence page
|  | And | choose the default price
|  | Then | the payment is in Euros
|  | And | UK VAT is not applied
ViewInvoiceEuro:
| Tags: | @FR031-3, @UR029-3
|  | Given | I have paid for a subscription for an EU user
|  | When | I access my billing portal
|  | Then | I can view an invoice for each month
|  | And | the invoice contains the price paid
|  | And | the invoice contains the customer name and address
|  | And | the invoice contains a statement regarding reverse tax

## CUSTOMERWORKFLOWUK
CustomerWorkflowUK
| Tags: | @FR, @FR030, @UR028, @Admin, @DesktopOnly
ShowRegistrationForm
| Tags: | @FR030-1, @UR028-1
|  | Given | no licence exists for the current machine
|  | When | I start Assistdent
|  | Then | the user registration form is displayed
CanRegisterInUK
| Tags: | @FR030-2, @UR028-1
|  | Given | no licence exists for the current machine
|  | And | I start Assistdent
|  | And | the use registration form is displayed
|  | When | I enter valid details
|  | And | I select "United Kingdom" for the country
|  | Then | the registration is successful
|  | And | AssistDent starts
TakePaymentUK
| Tags: | @FR030-3, @UR028-2
|  | Given | I have register a trial licence for a UK user
|  | When | I access the purchase licence page
|  | And | choose the default price
|  | Then | the payment is in GBP
|  | And | UK VAT is applied
PurchaseLicenseExtension
| Tags: | @FR030-4, @UR028-3
|  | When | I choose the Purchase Upgrade menu option
|  | And | I fill in my payment details
|  | And | I complete the purchase
|  | And | I restart the desktop app
|  | Then | the license period is extended by the amount specified for the purchased product
ViewInvoiceUK:
| Tags: | @FR030-5, @UR028-5
|  | Given | I have paid for a subscription for a UK user
|  | When | I access my billing portal
|  | Then | I can view an invoice for each month
|  | And | the invoice contains the price paid
|  | And | the invoice contains the customer name and address
|  | And | the invoice contains the VAT amount
|  | And | the invoice contains the MIL VAT number

## DELETEROI
DeleteROI
| Tags: | @FR, @FR006, @UR006, @Analysis
DisplayDeleteButton
| Tags: | @FR006-1, @UR006-1, @WebOnly
|  | Given | 2 submitted bitewing images
|  | Then | the images shall be displayed in tile view
|  | And | a delete button shall be displayed alongside each image
DeleteROI
| Tags: | @FR006-2, @UR006-1
|  | Given | an analyzed bitewing image which contains caries ROIs
|  | When | I select an ROI
|  | And | click the delete button corresponding to that image
|  | Then | the selected caries ROI will be deleted
DeleteROIs
| Tags: | @FR006-3, @UR006-1
|  | Given | an analyzed bitewing image which contains caries ROIs
|  | When | I select 2 or more ROIs
|  | And | click the delete button corresponding to that image
|  | Then | the selected caries ROIs will be deleted

## DESKTOPLICENSING
DesktopLicensing
| Tags: | @FR, @FR027, @UR025, @Admin, @DesktopOnly
CannotRegisterInNonCECountry
| Tags: | @FR027-1, @UR025-2
|  | Given | no licence exists for the current machine
|  | And | I start Assistdent
|  | And | the use registration form is displayed
|  | When | I enter valid details
|  | And | I select "United States" for the country
|  | Then | the registration is unsuccesful
|  | And | AssistDent does not start
RegisterUser
| Tags: | @FR027-2, @UR025-3
|  | Given | no licence exists for the current machine
|  | When | I start the desktop app
|  | Then | the user registration form is displayed
|  | When | I enter the registration details
|  | Then | the desktop app starts
ExpiredLicenceRestrictsAccess
| Tags: | @FR027-3, @UR025-5
|  | Given | a licence exists for the current machine
|  | And | the licence is expired
|  | When | I start the desktop app
|  | Then | a message tells me to upgrade
ExpiredLicenceRedirectsToUpgrade
| Tags: | @FR027-4, @UR025-5
|  | Given | a licence exists for the current machine
|  | And | the licence is expired
|  | When | I start the desktop app
|  | Then | the upgrade trial licence web page is displayed
ManualUpgradeAvailable
| Tags: | @FR027-5, @UR025-5
|  | Given | a licence exists for the current machine
|  | And | the licence is not expired
|  | When | I start the desktop app
|  | Then | the menu includes an option to upgrade
ManualUpgradeLinksToStore
| Tags: | @FR027-6, @UR025-5
|  | Given | a licence exists for the current machine
|  | And | the licence is not expired
|  | And | I have started the desktop app
|  | When | I choose the menu option "Upgrade Licence"
|  | Then | the upgrade trial licence web page is displayed
|  | And | the choice of upgrade options is approriate to my country
ValidLicence
| Tags: | @FR027-7, @UR025-6
|  | Given | a licence exists for the current machine
|  | And | the licence is not expired
|  | When | I start the desktop app
|  | Then | the desktop app opens
DisplayLicence
| Tags: | @FR027-8, @UR025-6
|  | Given | a licence exists for the current machine
|  | And | the licence is not expired
|  | When | I open the About Box
|  | Then | the licence key is displayed
|  | And | the licence expiry date is displayed
|  | And | the licensee is displayed
EnterNewLicenceKey
| Tags: | @FR027-9, @UR025-7
|  | Given | a licence exists for the current machine
|  | And | the licence is not expired
|  | And | I have started the desktop app
|  | When | I open the About Box
|  | And | I click Enter Key
|  | And | I enter a new valid key from the same CryptoLens product
|  | Then | the key is accepted
|  | When | I restart the application
|  | Then | the application restarts
EnterNewLicenceKeyNewProduct
| Tags: | @FR027-10, @UR025-7
|  | Given | a licence exists for the current machine
|  | And | the licence is not expired
|  | And | I have started the desktop app
|  | When | I open the About Box
|  | And | I click Enter Key
|  | And | I enter a new valid key prefixed with a new product id in the form "{product-id}:{licence-key}"
|  | Then | the key is accepted
|  | When | I restart the application
|  | Then | the application restarts
ConfigureTrialPeriod
| Tags: | @FR027-11, @UR025-8
|  | Given | no licence exists for the current machine
|  | When | I configure the store app to generate a <num1> day(s) licence
|  | And | I register for a trial licence
|  | Then | the licence expiry date can be seen in the application
|  | And | the expiry date is in <num2> day(s)
Examples
|  |  | num1 | num2
|  |  | 1 | 1
|  |  | 14 | 14
|  |  | 30 | 30

## DETECTIONLOGGING
DetectionLogging
| Tags: | @FR, @FR021, @UR019, @Admin, @WebOnly
ViewLoggedUserDetection
| Tags: | @FR021-1, @UR019-1
|  | Given | I am logged in with admin permissions
|  | And | I access <base>/analysis
|  | And | I submit 1 image and receive the results
|  | When | I access <base>/User
|  | And | click the Details link next to my username
|  | Then | I see a list of detections
|  | And | the list of detections includes an entry for the analysis performed
ViewLoggedAllowedIPDetection
| Tags: | @FR021-2, @UR019-2
|  | Given | I am not logged on
|  | And | an Allowed Ip exists for my computer
|  | And | I access <base>/analysis
|  | And | I submit 1 image and receive the results
|  | When | I log on with admin permissions
|  | And | I access <base>/AllowedIPs
|  | And | click the Details link next the the Allowed IP for my computer
|  | Then | I see a list of detections
|  | And | the list of detections includes an entry for the analysis performed

## DICOMIMAGEFORMAT
DICOMImageFormat
| Tags: | @FR, @FR017, @UR016, @Analysis, @WebOnly
CanAnalyseTiffImages
| Tags: | @FR017-1, @UR016-1, @UR016-4
|  | Given | I have logged in with analysis permissions
|  | When | I enter URL <base>/Analysis
|  | And | click the Select File area
|  | And | browse for a a bitewing image in DICOM format
|  | And | submit the image
|  | Then | the image is displayed correctly
|  | And | the results are displayed

## DISPLAYIMAGE
DisplayImage
| Tags: | @FR, @FR001, @UR001, @Analysis
The image shall be displayed on screen
DisplayImage
| Tags: | @FR001-1, @UR001-1
|  | Given | a bitewing image
|  | When | submitted for analysis
|  | Then | the bitewing image is displayed on screen
DisplayImages
| Tags: | @FR001-2, @UR001-2, @WebOnly
|  | Given | multiple bitewing images
|  | When | submitted for analysis
|  | Then | all bitewing images are displayed on screen

## FEATURETOUR
FeatureTour
| Tags: | @FR, @FR012, @UR012, @Labelling, @WebOnly
StartFeatureTour
| Tags: | @FR012-1, @UR012-1
|  | When | I access the tour URL
|  | Then | the tour starts
RetakeFeatureTour
| Tags: | @FR012-2, @UR012-1
|  | When | I access the tour URL a subsequent time
|  | Then | the tour starts
FollowFeatureTour
| Tags: | @FR012-3, @UR012-1
|  | Given | the tour is started
|  | When | I press Next
|  | Then | the tour moves to the next step
|  | And | I can repeatedly press next until the end of the tour is reached

## FOLDERWATCHER
FolderWatcher
| Tags: | @FR, @FR033, @UR030, @Admin, @DesktopOnly
StartsAutomaticallyAndRunsInSystemTray
| Tags: | @FR033-01, @UR030-02
|  | When | I log in to Windows
|  | Then | the Folder Watcher starts automatically
|  | And | an app appears in the system tray
AutomaticallyAnalyzes
| Tags: | @FR033-02, @UR030-01
|  | Given | the folder watcher is enabled
|  | When | I copy an image file to the watched folder
|  | Then | AssistDent automatically starts and analyzes the image
DisableAnalysis
| Tags: | @FR033-03, @UR030-03
|  | Given | the folder watcher is enabled
|  | When | I access the settings via the system tray app
|  | And | disable "automatic assist"
|  | And | I copy an image file to the watched folder
|  | Then | AssistDent does not start automatically
ConfigureWatchedFolder
| Tags: | @FR033-04, @UR030-04
|  | Given | the folder watcher is enabled
|  | When | the folder watcher is configured to watch a new folder
|  | And | I copy an image file to the previously watched folder
|  | Then | AssistDent does not start automatically
|  | When | I copy an image file to the watched folder
|  | Then | AssistDent automatically starts and analyzes the image
IgnoreMultipleFiles
| Tags: | @FR033-05, @UR030-01
|  | Given | the folder watcher is enabled
|  | When | I copy 2 or more image files at once to the folder
|  | Then | AssistDent does not start
|  | And | an informative message is displayed

## FULLSCREEN
FullScreen
| Tags: | @FR, @FR005, @UR005, @Analysis, @WebOnly
The image shall be displayed expanded to fill as much of the screen as possible
AccessFullScreen
| Tags: | @FR005-1, @UR005-1
|  | Given | 2 submitted bitewing images
|  | Then | the images shall be displayed in tile view
|  | And | an expand button shall be displayed alongside each image
EnterFullScreen
| Tags: | @FR005-2, @UR005-1
|  | Given | 2 images displayed in tile view
|  | When | I press the the expand button corresponding to the first image
|  | Then | the first image is displayed to fill as much of the screen as possible
|  | And | the other image is hidden from view
AccessTileView
| Tags: | @FR005-3, @UR005-1
|  | Given | an image displayed full screen
|  | Then | an exit/collapse button shall be displayed alongside the image
ReturnToTileView
| Tags: | @FR005-4, @UR005-1
|  | Given | an image displayed full screen
|  | When | I press the exit/collapse button the image is displayed in tile view

## GENERATEANDVIEWFEEDBACK
GenerateAndViewFeedback
| Tags: | @FR, @FR022, @UR020, @UR021, @Admin, @WebOnly
UserCanAccessFeedbackForm
| Tags: | @FR022-1, @UR021-1
|  | Given | I am logged on with analysis permissions
|  | When | I access the homepage
|  | Then | I can see the feedback link
|  | When | I access the analysis page
|  | Then | I can see the feedback link
|  | When | I follow the feedback link
|  | Then | I see the feedback page
|  | When | I access <base>/Feedback/New
|  | Then | I see the feedback page
UserCanEnterFeedback
| Tags: | @FR022-2, @UR021-1
|  | Given | I am logged in with analysis permissions
|  | And | I have accessed the feedback form
|  | When | I enter text into the feedback form
|  | And | click Submit Feedback
|  | Then | i am redirected to the homepage
AdminCanAccessUserFeedback
| Tags: | @FR022-3, @UR020-1
|  | Given | feedback has been entered for an identified user
|  | And | I am logged in with admin permissions
|  | And | I access <base>/User
|  | And | I click on the Details link next to the user
|  | Then | I can see a list of the user's feedback
|  | And | the list contains the date of the feedback
|  | And | the list contains the text of the feedback
AllowedIPCanAccessFeedbackForm
| Tags: | @FR022-4, @UR021-2
|  | Given | I am not logged on
|  | And | an Allowed IP exists which includes my computer's IP address
|  | When | I access the homepage
|  | Then | I can see the feedback link
|  | When | I access the analysis page
|  | Then | I can see the feedback link
|  | When | I follow the feedback link
|  | Then | I see the feedback page
|  | When | I access <base>/Feedback/New
|  | Then | I see the feedback page
AllowedIPCanEnterFeedback
| Tags: | @FR022-5, @UR021-2
|  | Given | I am not logged on
|  | And | an Allowed IP exists which includes my computer's IP address
|  | And | I have accessed the feedback form
|  | When | I enter text into the feedback form
|  | And | click Submit Feedback
|  | Then | i am redirected to the homepage
AdminCanAccessAllowedIPFeedback
| Tags: | @FR022-6, @UR020-2
|  | Given | feedback has been entered for an Allowed IP
|  | And | I am logged in with admin permissions
|  | And | I access <base>/AllowedIPs
|  | And | I click on the Details link next to the Allowed IP
|  | Then | I can see a list of the feedback associated from that IP
|  | And | the list contains the date of the feedback
|  | And | the list contains the text of the feedback

## IMAGEEXPORT
ImageExport
| Tags: | @FR, @FR034, @UR031, @Analysis, @DesktopOnly
ExportImageButton
| Tags: | @FR034-01, @UR031-01
|  | Given | I have performed caries detection
|  | When | I click the Export button
|  | Then | an image file is created
|  | And | that image contains the source X-Ray and annoted caries ROIs
ExportROIsAndTeeth
| Tags: | @FR034-02, @UR031-01
|  | Given | I have performed caries detection
|  | And | I configure the export to include teeth
|  | When | I click the Export button
|  | Then | an image file is created
|  | And | that image contains the source X-Ray and annoted caries ROIs, and teeth indicators
ExportROIsOnly
| Tags: | @FR034-03, @UR031-01
|  | Given | I have performed caries detection
|  | And | I configure the export to exclude teeth
|  | When | I click the Export button
|  | Then | an image file is created
|  | And | that image contains the source X-Ray and annoted caries ROIs but no teeth
ExportToNewLocation
| Tags: | @FR034-04, @UR031-01
|  | Given | I have performed caries detection
|  | And | I configure the export to be created in a new folder
|  | When | I click the Export button
|  | Then | an image file is created in the new location
|  | And | that image contains the source X-Ray and annoted caries ROIs
ExportHasUniqueName
| Tags: | @FR034-05, @UR031-01
|  | Given | I have performed caries detection
|  | When | I click the Export button
|  | Then | an image file is created
|  | And | that image file has a unique name based on the date and time
ExportCreatedOnExit
| Tags: | @FR034-06, @UR031-02
|  | Given | I have performed caries detection
|  | And | I have configured AssistDent to export the image on exit
|  | When | I close AssistDent
|  | Then | an image file is created
|  | And | that image contains the source X-Ray and annoted caries ROIs
SnapshotImage
| Tags: | @FR034-07, @UR031-03
|  | Given | I have performed caries detection
|  | When | I click the Snapshot button
|  | Then | an image file is copied to the clipboard
|  | And | that image contains the source X-Ray and annoted caries ROIs
RunCommandAfterExport
| Tags: | @FR034-08, @UR031-04
|  | Given | I have performed caries detection
|  | And | I have configured AssisDent to run a command after export
|  | When | I click the Export button
|  | Then | the export is created
|  | And | the configured command is executed
ExportToVistaSoft
| Tags: | @FR034-09, @UR031-04
|  | Given | I have performed caries detection
|  | And | I have configured AssisDent to send to VistaSoft after export
|  | When | I click the Export button
|  | Then | the export is created
|  | And | the image appears alongside the parent image in VistaSoft

## IMAGEID
ImageID
| Tags: | @FR, @FR025, @Analysis, @WebOnly
ShowImageIDFilename
| Tags: | @FR025-1, @UR023-1
|  | Given | I have submitted an image for analysis
|  | Then | the image ID is shown
|  | And | the image ID displays the image filename
ShowImageIDEditBox
| Tags: | @FR025-2, @UR023-2
|  | Given | I have submitted an image for analysis
|  | And | the image ID displays the image filename
|  | When | I hover over the ID
|  | Then | a box is shown indicating that the field can be edited

## IMAGEMAGNIFER
ImageMagnifer
| Tags: | @FR, @FR004, @UR004, @Analysis
With the mouse button held down, the section of the image directly under the mouse pointer will be zoomed to 3x magnification
ShowMagnifer
| Tags: | @FR004-1, @UR004-1
|  | Given | a bitewing image displayed on screen
|  | When | I hover the mouse pointer over a location in the image
|  | And | I hold the mouse button down
|  | Then | the magnifier should be displayed
|  | And | the image location should be displayed at 3x magnification
MoveMagnifer
| Tags: | @FR004-2, @UR004-1
|  | Given | a bitewing image displayed on screen
|  | And | the magnifier is currently displayed
|  | When | I move the mouse pointer to another location in the image
|  | Then | the image location should be updated
HideAnnotationInMagnifer
| Tags: | @FR004-3, @UR004-2
|  | Given | a bitewing image displayed on screen
|  | And | the magnifier is currently displayed
|  | And | the magnifer is positioned over an annotation
|  | Then | no annotation should be displayed in the magnification

## IMAGESUBMIT
ImageSubmit
| Tags: | @FR, @FR018, @UR016, @Analysis, @WebOnly
CanSubmitImageByBrowsing
| Tags: | @FR018-1, @UR016-1
|  | Given | I have logged in with analysis permissions
|  | When | I enter URL <base>/Analysis
|  | And | click the dashed submit area
|  | And | browse for a bitewing image
|  | And | submit the image
|  | Then | the image is displayed correctly
|  | And | the results are displayed
CanSubmitImageByDragAndDrop
| Tags: | @FR018-1, @UR016-2
|  | Given | I have logged in with analysis permissions
|  | When | I enter URL <base>/Analysis
|  | And | drag and drop a bitewing image from Explorer to the dashed submit area
|  | Then | the image is displayed correctly
|  | And | the results are displayed
CanSubmitImagesByBrowsing
| Tags: | @FR018-3, @UR016-1
|  | Given | I have logged in with analysis permissions
|  | When | I enter URL <base>/Analysis
|  | And | click the dashed submit area
|  | And | browse for 2 bitewing images
|  | And | submit the images
|  | Then | the images are displayed correctly
|  | And | the results are displayed
CanSubmitImagesByDragAndDrop
| Tags: | @FR018-4, @UR016-2
|  | Given | I have logged in with analysis permissions
|  | When | I enter URL <base>/Analysis
|  | And | drag and drop 2 bitewing images from Explorer to the dashed submit area
|  | Then | the images are displayed correctly
|  | And | the results are displayed

## LEGALINFORMATIONPAGE
LegalInformationPage
| Tags: | @FR, @FR015, @UR015, @Labelling
CanAccessLegalPage
| Tags: | @FR015-1, @UR015-1, @WebOnly
|  | When | I enter URL <base>/Legal
|  | Then | a page of legal information is displayed
|  | And | it contains the Regulatory clinical use information
CanAccessClinicalUse
| Tags: | @FR015-2, @UR015-1, @DesktopOnly
|  | When | I open the About menu option
|  | Then | a tab of clinical use is displayed
|  | And | it contains the clinical use information

## MANAGEALLOWEDIPRANGES
ManageAllowedIPRanges
| Tags: | @FR, @FR014, @UR014, @Security, @WebOnly
CanAccessAllowedIPRangeURL
| Tags: | @FR014-1, @UR014-1
|  | Given | I have logged in with admin permissions
|  | When | I enter URL <base>/AllowedIPs
|  | Then | a list of Allowed IP Ranges is displayed
|  | And | the name of each Allowed IP Range is listed
|  | And | the start address of each Allowed IP Range is listed
|  | And | the end address of each Allowed IP Range is listed
|  | And | the expiry date of each Allowed IP Range is listed
CanAddAllowedIPRange
| Tags: | @FR014-2, @UR014-2
|  | Given | I have logged in with admin permissions
|  | And | have accessed the list of Allowed IP Ranges
|  | When | I click the Allow IP Range button
|  | And | enter a name
|  | And | enter a start address
|  | And | enter an end address
|  | And | enter an expiry date
|  | And | click Add
|  | Then | the page is redirected to the list of Allowed IP Ranges
|  | And | the list of Allowed IP Ranges now includes the new entry
|  | And | the details match those entered in previous steps
CantAddInvalidAllowedIPRange
| Tags: | @FR014-3, @UR014-2
|  | Given | I have logged in with admin permissions
|  | And | have accessed the list of Allowed IP Ranges
|  | When | I click the Allow IP Range button
|  | And | enter a name
|  | And | enter a start address
|  | And | enter an end address
|  | And | enter the name again in the expiry date field
|  | And | click Add
|  | Then | a validation error message appears adjacent to the expiry date field
|  | And | the Allowed IP Range cannot be saved
CanEditAllowedIPRange
| Tags: | @FR014-4, @UR014-3
|  | Given | I have logged in with admin permissions
|  | And | I have created at least one Allowed IP Range
|  | When | I click the Edit link
|  | And | edit the name
|  | And | click the Save button
|  | Then | the name is updated in the list of Allowed IP Ranges
CantEditInvalidAllowedIPRange
| Tags: | @FR014-5, @UR014-3
|  | Given | I have logged in with admin permissions
|  | And | I have created at least one Allowed IP Range
|  | When | I click the Edit link
|  | And | repeat the name in the Expiry Date field
|  | And | click the Save button
|  | Then | a validation error message appears adjacent to the expiry date field
|  | And | the Allowed IP Range cannot be saved
CanDeleteAllowedIPRange
| Tags: | @FR014-6, @UR014-4
|  | Given | I have logged in with admin permissions
|  | And | I have created at least one Allowed IP Range
|  | When | I click the Delete link
|  | Then | the list of Allowed IP Ranges no longer includes the deleted record

## MANAGEUSERLOGINS
ManageUserLogins
| Tags: | @FR, @FR019, @UR018, @Security, @WebOnly
CanAccessManageUsersURL
| Tags: | @FR019-1, @UR018-1
|  | Given | I have logged in with admin permissions
|  | When | I enter URL <base>/User
|  | Then | a list of Users is displayed
|  | And | the name of each User is displayed
|  | And | the email address of each User is displayed
|  | And | the number of performed detections is displayed
CanAddUser
| Tags: | @FR019-2, @UR018-2
|  | Given | I have logged in with admin permissions
|  | And | have accessed the list of Users
|  | When | I click the New User button
|  | And | enter an email
|  | And | enter a full name
|  | And | enter comments
|  | And | enter a password
|  | And | confirm a password
|  | And | click Register
|  | Then | the page is redirected to the list of Users
|  | And | the list of Users now includes the new entry
|  | And | the details match those entered in previous steps
CantAddInvalidUser
| Tags: | @FR019-3, @UR018-2
|  | Given | I have logged in with admin permissions
|  | And | have accessed the list of Users
|  | When | I click the New User button
|  | And | click Register
|  | Then | the validation error for email address is displayed
|  | And | the validation error for password is displayed
|  | When | I enter "name" in the email field
|  | And | click Register
|  | Then | the validation error a valid email address is displayed
|  | And | the validation error for password is displayed
|  | When | I enter "name@email.com" in the email field
|  | And | click Register
|  | Then | the validation error for password is displayed
|  | When | I enter "12345" in the password field
|  | And | click Register
|  | Then | the validation error for password complexity is displayed
|  | When | I enter "123456" in the password field
|  | And | click Register
|  | Then | the validation error for password confirmation is displayed
|  | When | I enter "notamatch" in the password field
|  | And | click Register
|  | Then | the validation error for password confirmation is displayed
CanDeleteAllowedIPRange
| Tags: | @FR019-4, @UR018-3
|  | Given | I have logged in with admin permissions
|  | And | I have created at least one User
|  | When | I click the Delete link
|  | Then | the list of Users no longer includes the deleted record

## PROGRESSDISPLAY
ProgressDisplay
| Tags: | @FR, @FR028, @UR026, @Analysis, @DesktopOnly
ShowAnalysisFeedback
| Tags: | @FR028-1, @UR026-1
|  | When | I submit an image for analysis
|  | Then | the app displays status updates
|  | And | the updates are clearly visible

## REPOSITIONROI
RepositionROI
| Tags: | @FR, @FR009, @UR009, @Analysis
MoveAndDragROI
| Tags: | @FR009-1, @UR009-1
|  | Given | an analyzed bitewing image which contains caries ROIs
|  | When | I click and drag an ROI
|  | Then | the ROI will move with the mouse
|  | And | the ROI selection status will remain unchanged
MoveROI
| Tags: | @FR009-2, @UR009-1
|  | Given | an analyzed bitewing image which contains caries ROIs
|  | When | I click and drag an ROI
|  | And | move the mouse
|  | And | release the mouse
|  | Then | the ROI will stay at its new location
|  | And | the ROI selection status will remain unchanged
FlipROI
| Tags: | @FR009-3, @UR009-2
|  | Given | an analyzed bitewing image which contains caries ROIs
|  | When | I select an ROI
|  | And | click the flip button corresponding to that image
|  | Then | the selected caries ROI will be flipped
|  | And | the ROI will remain selected
FlipROIs
| Tags: | @FR009-4, @UR009-2
|  | Given | an analyzed bitewing image which contains caries ROIs
|  | When | I select 2 or more ROIs
|  | And | click the flip button corresponding to that image
|  | Then | the selected caries ROIs will be flipped
|  | And | the ROIs will remain selected

## ROIINDICATOR
ROIIndicator
| Tags: | @FR, @FR003, @UR003, @Analysis
The location of caries regions of interest shall be displayed in position on the image
ROIIndicators
| Tags: | @FR003-1, @UR003-1
|  | Given | a bitewing image known to contain caries
|  | When | caries detection is complete
|  | Then | caries ROIs are indicated over the image
ROIArrows
| Tags: | @FR003-2, @UR003-1
|  | Given | a bitewing image known to contain caries
|  | When | caries detection is complete
|  | Then | caries ROIs are indicated by arrows
|  | And | the arrows do not obscure the tooth edge
ROIEasilySeen
| Tags: | @FR003-3, @UR003-1
|  | Given | a bitewing image known to contain caries
|  | When | caries detection is complete
|  | Then | caries ROI indicators are a consistent size
|  | And | caries ROI indicators are brightly coloured
ROIConsistentSize
| Tags: | @FR003-4, @UR003-1, @WebOnly
|  | Given | 2 bitewing images of different resolution
|  | When | caries detection is complete
|  | Then | caries ROI indicators in both images are a consistent size

## SEEWHATSNEW
SeeWhatsNew
| Tags: | @FR, @FR029, @UR027, @Analysis, @DesktopOnly
SeeWhatsNewOnStartup
| Tags: | @FR029-1, @UR027-1
|  | Given | I have not dismissed the news for the version installed
|  | When | I start AssistDent
|  | Then | the what's new form is displayed
SeeWhatsNewAllVersions
| Tags: | @FR029-2, @UR027-1
|  | Given | an upgrade of AssistDent from the previous version
|  | And | the new features were previously dismissed
|  | When | I start AssistDent
|  | Then | the what's new form is displayed
|  | And | the what's new form contains the new features
|  | And | the new features apply only to the current version
SeeWhatsNewAllVersions
| Tags: | @FR029-3, @UR027-1
|  | Given | a clean install of AssistDent with no settings files
|  | When | I start AssistDent
|  | Then | the what's new form is displayed
|  | And | the what's new form contains the new features
|  | And | the new features apply to the current version and every previous version
DimissAndDontShowAgain
| Tags: | @FR029-4, @UR027-2
|  | Given | I have not dismissed the news for the version installed
|  | And | I start AssistDent
|  | And | the what's new form is displayed
|  | When | I check the "Don't show me this again" checkbox
|  | And | restart AssistDent
|  | Then | AssistDent starts
|  | And | the what's new form is not displayed
SeeWhatsNewAfterStartup
| Tags: | @FR029-5, @UR027-3
|  | Given | I have started AssistDent
|  | When | I access the "See What' New" menu option
|  | Then | the what's new form is displayed

## SHELLINTEGRATION
ShellIntegration
| Tags: | @FR, @FR035, @UR033, @Admin, @DesktopOnly
OpenWith
| Tags: | @FR035-01, @UR033-01
|  | Given | I have an image file
|  | When | in Explorer I right click choose open with and find AssistDent.exe
|  | Then | the image opens in AssistDent
DontOpenNonImage
| Tags: | @FR035-02, @UR033-01
|  | Given | I have a text file
|  | When | in Explorer I right click choose open with and find AssistDent.exe
|  | Then | nothing happens
DontOpenMultipleImages
| Tags: | @FR035-03, @UR033-01
|  | Given | I have multiple image files
|  | When | in Explorer I right click choose open with and find AssistDent.exe
|  | Then | nothing happens
OpenWith
| Tags: | @FR035-04, @UR033-02
|  | Given | I have an image file
|  | When | in Explorer I right click choose Open With AssistDent
|  | Then | the image opens in AssistDent
DontOpenNonImage
| Tags: | @FR035-05, @UR033-02
|  | Given | I have a text file
|  | When | in Explorer I right click
|  | Then | the Open With AssistDent menu option is not visible
DontOpenMultipleImages
| Tags: | @FR035-06, @UR033-02
|  | Given | I have multiple image files
|  | When | in Explorer I right click
|  | Then | the Open With AssistDent menu option is not visible

## TIFFIMAGEFORMAT
TiffImageFormat
| Tags: | @FR, @FR016, @UR016, @Analysis, @WebOnly
CanAnalyseTiffImages
| Tags: | @FR016-1, @UR016-1, @UR016-3
|  | Given | I have logged in with analysis permissions
|  | When | I enter URL <base>/Analysis
|  | And | click the Select File area
|  | And | browse for a a bitewing image in Tiff format
|  | And | submit the image
|  | Then | the image is displayed correctly
|  | And | the results are displayed

## TOGGLEANNOTATION
ToggleAnnotation
| Tags: | @FR, @FR007, @UR007, @Analysis
AccessAnnotationToggle
| Tags: | @FR007-1, @UR007-1, @WebOnly
|  | Given | 2 submitted bitewing images
|  | Then | the images shall be displayed in tile view
|  | And | a hide annotation button shall be displayed alongside each image
HideImageAnnotation
| Tags: | @FR007-2, @UR007-1
|  | Given | an analyzed bitewing image which contains caries ROIs
|  | When | I click the corresponding hide annotation button
|  | Then | the tooth indicators are hidden
|  | And | the caries ROI indicators are hidden
|  | And | the hide annotation button is replaced with a show annotation button
RestoreImageAnnotation
| Tags: | @FR007-3, @UR007-2
|  | Given | an analyzed bitewing image which contains caries ROIs
|  | And | the annotation is currently hidden
|  | When | I click the corresponding show annotation button
|  | Then | the tooth indicators are displayed
|  | And | the caries ROI indicators are displayed
|  | And | the show annotation button is replaced with a hide annotation button

## TOOTHINDICATOR
ToothIndicator
| Tags: | @FR, @FR002, @UR002, @Analysis
Detected teeth shall be displayed in position over the image
ToothIndicators
| Tags: | @FR002-1, @UR002-1
|  | Given | a bitewing image known to contain teeth
|  | When | caries detection is complete
|  | Then | the detected teeth are displayed on the image

## UPDATECLIENT
UpdateClient
| Tags: | @FR, @FR032, @UR024, @Admin, @DesktopOnly
UpdateNotAvailable
| Tags: | @FR032-1, @UR024-1
|  | Given | an update is not available
|  | When | I start the app
|  | Then | the app starts
UpdateAvailable
| Tags: | @FR032-2, @UR024-2, @UR024-3
|  | Given | an update is available
|  | When | I start the app
|  | Then | the app starts
|  | And | a message is displayed
|  | And | the message indicates the availability of an update
|  | And | the message includes a link

## CARIESDETECTIONPERFORMANCE
CariesDetectionPerformance
| Tags: | @NFR, @NFR002, @UR003, @Performance, @WebOnly
Measure the performance on the evaluation image set
PerformanceEvaluation
| Tags: | @NFR002-1, @UR003
|  | Given | I am logged in with analysis permissions
|  | When | I evaluate the images
|  |  | image | goldstandard
|  |  | TestData\Images\SC22.tif | TestData\GoldStandard\SC22.tif.caries.json
|  |  | TestData\Images\SC28.tif | TestData\GoldStandard\SC28.tif.caries.json
|  |  | TestData\Images\SC56.tif | TestData\GoldStandard\SC56.tif.caries.json
|  |  | TestData\Images\SC64.tif | TestData\GoldStandard\SC64.tif.caries.json
|  |  | TestData\Images\SC75.tif | TestData\GoldStandard\SC75.tif.caries.json
|  |  | TestData\Images\SC85.tif | TestData\GoldStandard\SC85.tif.caries.json
|  |  | TestData\Images\SC94.tif | TestData\GoldStandard\SC94.tif.caries.json
|  |  | TestData\Images\ML29.tif | TestData\GoldStandard\ML29.tif.caries.json
|  |  | TestData\Images\ML45.tif | TestData\GoldStandard\ML45.tif.caries.json
|  |  | TestData\Images\ML51.tif | TestData\GoldStandard\ML51.tif.caries.json
|  |  | TestData\Images\ML83.tif | TestData\GoldStandard\ML83.tif.caries.json
|  |  | TestData\Images\ML89.tif | TestData\GoldStandard\ML89.tif.caries.json
|  |  | TestData\Images\ML95.tif | TestData\GoldStandard\ML95.tif.caries.json
|  |  | TestData\Images\ML102.tif | TestData\GoldStandard\ML102.tif.caries.json
|  |  | TestData\Images\ML123.tif | TestData\GoldStandard\ML123.tif.caries.json
|  |  | TestData\Images\ML128.tif | TestData\GoldStandard\ML128.tif.caries.json
|  |  | TestData\Images\ML135.tif | TestData\GoldStandard\ML135.tif.caries.json
|  |  | TestData\Images\RC8.tif | TestData\GoldStandard\RC8.tif.caries.json
|  |  | TestData\Images\RC36.tif | TestData\GoldStandard\RC36.tif.caries.json
|  |  | TestData\Images\RC43.tif | TestData\GoldStandard\RC43.tif.caries.json
|  |  | TestData\Images\RC51.tif | TestData\GoldStandard\RC51.tif.caries.json
|  |  | TestData\Images\RC56.tif | TestData\GoldStandard\RC56.tif.caries.json
|  |  | TestData\Images\RC64.tif | TestData\GoldStandard\RC64.tif.caries.json
|  |  | TestData\Images\RC79.tif | TestData\GoldStandard\RC79.tif.caries.json
|  |  | TestData\Images\RC104.tif | TestData\GoldStandard\RC104.tif.caries.json
|  |  | TestData\Images\RC108.tif | TestData\GoldStandard\RC108.tif.caries.json
|  |  | TestData\Images\RC133.tif | TestData\GoldStandard\RC133.tif.caries.json
|  |  | TestData\Images\RC136.tif | TestData\GoldStandard\RC136.tif.caries.json
|  |  | TestData\Images\TS4.tif | TestData\GoldStandard\TS4.tif.caries.json
|  |  | TestData\Images\TS19.tif | TestData\GoldStandard\TS19.tif.caries.json
|  |  | TestData\Images\TS21.tif | TestData\GoldStandard\TS21.tif.caries.json
|  |  | TestData\Images\TS45.tif | TestData\GoldStandard\TS45.tif.caries.json
|  |  | TestData\Images\TS50.tif | TestData\GoldStandard\TS50.tif.caries.json
|  |  | TestData\Images\TS75.tif | TestData\GoldStandard\TS75.tif.caries.json
|  |  | TestData\Images\TS91.tif | TestData\GoldStandard\TS91.tif.caries.json
|  |  | TestData\Images\TS95.tif | TestData\GoldStandard\TS95.tif.caries.json
|  |  | TestData\Images\TS129.tif | TestData\GoldStandard\TS129.tif.caries.json
|  |  | TestData\Images\TS146.tif | TestData\GoldStandard\TS146.tif.caries.json
|  |  | TestData\Images\TS152.tif | TestData\GoldStandard\TS152.tif.caries.json
|  |  | TestData\Images\TS212.tif | TestData\GoldStandard\TS212.tif.caries.json
|  |  | TestData\Images\TS213.tif | TestData\GoldStandard\TS213.tif.caries.json
|  |  | TestData\Images\TS217.tif | TestData\GoldStandard\TS217.tif.caries.json
|  |  | TestData\Images\TS223.tif | TestData\GoldStandard\TS223.tif.caries.json
|  |  | TestData\Images\TS230.tif | TestData\GoldStandard\TS230.tif.caries.json
|  |  | TestData\Images\TS236.tif | TestData\GoldStandard\TS236.tif.caries.json
|  |  | TestData\Images\TS237.tif | TestData\GoldStandard\TS237.tif.caries.json
|  |  | TestData\Images\TS296.tif | TestData\GoldStandard\TS296.tif.caries.json
|  |  | TestData\Images\TS299.tif | TestData\GoldStandard\TS299.tif.caries.json
|  |  | TestData\Images\TS305.tif | TestData\GoldStandard\TS305.tif.caries.json
|  |  | TestData\Images\TS313.tif | TestData\GoldStandard\TS313.tif.caries.json
|  |  | TestData\Images\JD16.tif | TestData\GoldStandard\JD16.tif.caries.json
|  |  | TestData\Images\JD43.tif | TestData\GoldStandard\JD43.tif.caries.json
|  |  | TestData\Images\JD73.tif | TestData\GoldStandard\JD73.tif.caries.json
|  |  | TestData\Images\JD91.tif | TestData\GoldStandard\JD91.tif.caries.json
|  |  | TestData\Images\JD98.tif | TestData\GoldStandard\JD98.tif.caries.json
|  |  | TestData\Images\JD107.tif | TestData\GoldStandard\JD107.tif.caries.json
|  |  | TestData\Images\JD131.tif | TestData\GoldStandard\JD131.tif.caries.json
|  |  | TestData\Images\JD142.tif | TestData\GoldStandard\JD142.tif.caries.json
|  |  | TestData\Images\JD143.tif | TestData\GoldStandard\JD143.tif.caries.json
|  |  | TestData\Images\CR26.dcm | TestData\GoldStandard\CR26.dcm.caries.json
|  |  | TestData\Images\CR33.dcm | TestData\GoldStandard\CR33.dcm.caries.json
|  |  | TestData\Images\CR44.dcm | TestData\GoldStandard\CR44.dcm.caries.json
|  |  | TestData\Images\CR83.dcm | TestData\GoldStandard\CR83.dcm.caries.json
|  |  | TestData\Images\CR98.dcm | TestData\GoldStandard\CR98.dcm.caries.json
|  |  | TestData\Images\CR108.dcm | TestData\GoldStandard\CR108.dcm.caries.json
|  |  | TestData\Images\CR117.dcm | TestData\GoldStandard\CR117.dcm.caries.json
|  |  | TestData\Images\CR128.dcm | TestData\GoldStandard\CR128.dcm.caries.json
|  |  | TestData\Images\CR138.dcm | TestData\GoldStandard\CR138.dcm.caries.json
|  |  | TestData\Images\CR139.dcm | TestData\GoldStandard\CR139.dcm.caries.json
|  |  | TestData\Images\CR145.dcm | TestData\GoldStandard\CR145.dcm.caries.json
|  |  | TestData\Images\VC11.dcm | TestData\GoldStandard\VC11.dcm.caries.json
|  |  | TestData\Images\VC26.dcm | TestData\GoldStandard\VC26.dcm.caries.json
|  |  | TestData\Images\VC28.dcm | TestData\GoldStandard\VC28.dcm.caries.json
|  |  | TestData\Images\VC35.dcm | TestData\GoldStandard\VC35.dcm.caries.json
|  |  | TestData\Images\VC45.dcm | TestData\GoldStandard\VC45.dcm.caries.json
|  |  | TestData\Images\VC60.dcm | TestData\GoldStandard\VC60.dcm.caries.json
|  |  | TestData\Images\VC87.dcm | TestData\GoldStandard\VC87.dcm.caries.json
|  |  | TestData\Images\VC99.dcm | TestData\GoldStandard\VC99.dcm.caries.json
|  |  | TestData\Images\VC106.dcm | TestData\GoldStandard\VC106.dcm.caries.json
|  |  | TestData\Images\VC113.dcm | TestData\GoldStandard\VC113.dcm.caries.json
|  |  | TestData\Images\AK2.dcm | TestData\GoldStandard\AK2.dcm.caries.json
|  |  | TestData\Images\AK48.dcm | TestData\GoldStandard\AK48.dcm.caries.json
|  |  | TestData\Images\AK51.dcm | TestData\GoldStandard\AK51.dcm.caries.json
|  |  | TestData\Images\AK77.dcm | TestData\GoldStandard\AK77.dcm.caries.json
|  |  | TestData\Images\AK82.dcm | TestData\GoldStandard\AK82.dcm.caries.json
|  |  | TestData\Images\AK84.dcm | TestData\GoldStandard\AK84.dcm.caries.json
|  |  | TestData\Images\AK88.dcm | TestData\GoldStandard\AK88.dcm.caries.json
|  |  | TestData\Images\AK89.dcm | TestData\GoldStandard\AK89.dcm.caries.json
|  |  | TestData\Images\AK102.dcm | TestData\GoldStandard\AK102.dcm.caries.json
|  |  | TestData\Images\image21.tif | TestData\GoldStandard\image21.tif.caries.json
|  |  | TestData\Images\image26.tif | TestData\GoldStandard\image26.tif.caries.json
|  |  | TestData\Images\image33.tif | TestData\GoldStandard\image33.tif.caries.json
|  |  | TestData\Images\image37.tif | TestData\GoldStandard\image37.tif.caries.json
|  |  | TestData\Images\image55.tif | TestData\GoldStandard\image55.tif.caries.json
|  |  | TestData\Images\image56.tif | TestData\GoldStandard\image56.tif.caries.json
|  |  | TestData\Images\image60.tif | TestData\GoldStandard\image60.tif.caries.json
|  |  | TestData\Images\image84.tif | TestData\GoldStandard\image84.tif.caries.json
|  |  | TestData\Images\image105.tif | TestData\GoldStandard\image105.tif.caries.json
|  |  | TestData\Images\image127.tif | TestData\GoldStandard\image127.tif.caries.json
|  | And | I compute the {SensitivityD1D2} as sum(TPD1D2) / sum(GSD1D2)
|  | And | I compute the {FPPerImage} as mean(FP)
|  | And | I compute the {Precision} as sum(TP) / sum(Detections)
|  | Then | the {SensitivityD1D2} is at least 0.5
PerformanceEvaluationStudyImageSet
| Tags: | @NFR002-2, @UR003
|  | Given | I am logged in with analysis permissions
|  | When | I evaluate the images
|  |  | image | goldstandard
|  |  | TestData\Images\AK2.dcm | TestData\GoldStandard\AK2.dcm.caries.json
|  |  | TestData\Images\AK51.dcm | TestData\GoldStandard\AK51.dcm.caries.json
|  |  | TestData\Images\AK82.dcm | TestData\GoldStandard\AK82.dcm.caries.json
|  |  | TestData\Images\CR44.dcm | TestData\GoldStandard\CR44.dcm.caries.json
|  |  | TestData\Images\CR83.dcm | TestData\GoldStandard\CR83.dcm.caries.json
|  |  | TestData\Images\CR108.dcm | TestData\GoldStandard\CR108.dcm.caries.json
|  |  | TestData\Images\image33.tif | TestData\GoldStandard\image33.tif.caries.json
|  |  | TestData\Images\image37.tif | TestData\GoldStandard\image37.tif.caries.json
|  |  | TestData\Images\image56.tif | TestData\GoldStandard\image56.tif.caries.json
|  |  | TestData\Images\image84.tif | TestData\GoldStandard\image84.tif.caries.json
|  |  | TestData\Images\image127.tif | TestData\GoldStandard\image127.tif.caries.json
|  |  | TestData\Images\JD16.tif | TestData\GoldStandard\JD16.tif.caries.json
|  |  | TestData\Images\ML29.tif | TestData\GoldStandard\ML29.tif.caries.json
|  |  | TestData\Images\ML51.tif | TestData\GoldStandard\ML51.tif.caries.json
|  |  | TestData\Images\ML83.tif | TestData\GoldStandard\ML83.tif.caries.json
|  |  | TestData\Images\RC36.tif | TestData\GoldStandard\RC36.tif.caries.json
|  |  | TestData\Images\RC51.tif | TestData\GoldStandard\RC51.tif.caries.json
|  |  | TestData\Images\TS91.tif | TestData\GoldStandard\TS91.tif.caries.json
|  |  | TestData\Images\TS146.tif | TestData\GoldStandard\TS146.tif.caries.json
|  |  | TestData\Images\TS230.tif | TestData\GoldStandard\TS230.tif.caries.json
|  |  | TestData\Images\TS296.tif | TestData\GoldStandard\TS296.tif.caries.json
|  |  | TestData\Images\TS305.tif | TestData\GoldStandard\TS305.tif.caries.json
|  |  | TestData\Images\VC11.dcm | TestData\GoldStandard\VC11.dcm.caries.json
|  |  | TestData\Images\VC87.dcm | TestData\GoldStandard\VC87.dcm.caries.json
|  | And | I compute the {SensitivityD1D2} as sum(TPD1D2) / sum(GSD1D2)
|  | And | I compute the {FPPerImage} as mean(FP)
|  | And | I compute the {Precision} as sum(TP) / sum(Detections)
|  | Then | the {SensitivityD1D2} is at least 0.5

## ENCRYPTTRAFFIC
EncryptTraffic
| Tags: | @NFR, @NFR001, @UR011, @Security, @WebOnly
WebsiteUsesSSL
| Tags: | @NFR001-1, @UR011-2
|  | Given | I access the Assistdent web site
|  | When | I view the site security information
|  | Then | the security indicator reveals SSL
AllPagesSecured
| Tags: | @NFR001-2, @UR011-2
|  | Given | I enable network traffic view in developer tools
|  | When | I access each main page of the Assistdent web site
|  | Then | the request are submitted using https
|  | And | the responses returned are received using https
AnalysisSecured
| Tags: | @NFR001-3, @UR011-2
|  | Given | I enable network traffic view in developer tools
|  | When | I submit an image
|  | Then | the request is submitted using https
|  | And | the response returned is received using https

## USERDATA
UserData
| Tags: | @NFR, @NFR003, @UR018, @Security
CustomerDataStoredSecurely
| Tags: | @NFR003-1, @DesktopOnly
|  | Given | I have customers
|  | Then | other than on the local computer, customer details are storred securely and exclusively on CryptoLens and Stripe

## USERD(1)
UserData
| Tags: | @NFR, @NFR004, @UR032, @Admin, @DesktopOnly
RecordNumberOfUses
| Tags: | @NFR004-1, @UR032-1
|  | Given | AssistDent has been installed
|  | And | registered and used at least once
|  | When | I examine the data objects of any given user's licence
|  | Then | I can see the number of uses is recorded for that user
RecordDateLastUsed
| Tags: | @NFR004-2, @UR032-2
|  | Given | AssistDent has been installed
|  | And | registered and used at least once
|  | When | I examine the data objects of any given user's licence
|  | Then | I can see the date on which the user last used AssistDent

## ADDCARIESROI
AddCariesROI
| Tags: | @UR, @UR008, @Analysis
As a dentist I want to add caries regions of interest at a location I deem dubious So that I can record for future use and share the results with a patient
AddCariesROI
| Tags: | @UR008-1
|  | Given | I submit an image
|  | When | I choose a location at which to add a caries ROI to the image
|  | Then | the new caries ROI is show on the image

## ALLOWIPRANGE
AllowIPRange
| Tags: | @UR, @UR014, @Security, @WebOnly
As an administrator I want to create, edit and delete whitelisted or Allowed IP Ranges So that I can allow users accessing from those IP addresses to access analysis functionality without providing credentials
ListAllowedIPRange
| Tags: | @UR014-1
|  | Given | I am logged in with admin permissions
|  | When | I request the Allowed IP Ranges page
|  | Then | a list of allowed ip ranges is displayed
AddAllowedIPRange
| Tags: | @UR014-2
|  | Given | I am logged in with admin permissions
|  | When | I create an Allowed IP Range for my computer
|  | Then | my IP address is included in the list of Allowed IP Ranges
EditAllowedIPRange
| Tags: | @UR014-3
|  | Given | I am logged in with admin permissions
|  | And | a single Allowed IP Range exists that includes my computer or subdomain's IP address
|  | When | I edit the Allowed IP Range for my computer and enter a different computer's IP address
|  | Then | the IP address is updated in the list of Allowed IP Ranges
DeleteAllowedIPRange
| Tags: | @UR014-4
|  | Given | I am logged in with admin permissions
|  | And | a single Allowed IP Range exists that includes my computer or subdomain's IP address
|  | When | I delete the Allowed IP Range for my computer
|  | Then | the IP address is no longer present in the list of Allowed IP Ranges
ExpireAllowedIPRange
| Tags: | @UR-14-5
|  | Given | I am logged in with admin permissions
|  | And | a single Allowed IP Range exists that includes my computer or subdomain's IP address
|  | When | I set the Expiry Date to be some date in the past
|  | Then | the expiry date is updated in the list of Allowed IP Ranges

## BYPASSLO(1)
BypassLogin
| Tags: | @UR, @UR013, @Security, @WebOnly
As a dentist at an institution I want to bypass the standard login security So that I can access functionality immediately And so that I can access functionality from any computer at the institution without entering credentials
ConnectFromAllowedIP
| Tags: | @UR013-1
|  | Given | I am using a computer within an allowed IP range
|  | And | I am logged out of the application
|  | When | I request the analysis page
|  | Then | the analysis page is displayed
|  | When | I submit an image
|  | Then | the image and results are displayed
ConnectFromOtherIP
| Tags: | @UR013-2
|  | Given | I am using a computer that isn't included within an allowed IP range
|  | And | I am logged out of the application
|  | When | I request the analysis page
|  | Then | the login page is displayed
ConnectFromExpiredIP
| Tags: | @UR013-3
|  | Given | I am using a computer within an allowed IP range but an expiry date in the past
|  | And | I am logged out of the application
|  | When | I request the analysis page
|  | Then | the login page is displayed

## COMMANDLINEINTERF(1)
CommandLineInterface
| Tags: | @UR, @UR024, @Integration, @DesktopOnly
As a third-party developer I want to be able to pass command line arguments to AssistDent So that I can instruct it to analyze specific images And display custom info
SubmitImage
| Tags: | @UR024-1
|  | When | I submit an image filename to AssistDent from a third-party
|  | Then | AssistDent analyses the specified image
|  | And | displays the correct patient name

## CONTROBUTTONTOOLTIPS
ControButtonToolTips
| Tags: | @UR, @UR010, @Labelling, @WebOnly
As a dentist I want to be given indications of what the control buttons means so that I can understand the functionality
ControButtonToolTips
| Tags: | @UR010-1
|  | Given | I have submitted an image
|  | When | I hover over the results buttons below the image
|  | Then | tool-tips appear to indicate the function of each button

## CUSTOMERWORKFLO(1)
CustomerWorkflowEU
| Tags: | @UR, @UR029, @DesktopOnly, @Admin
As a company We want to register EU users and accept payment for subscriptions to AssistDent So we can make money
CanUseInEU
| Tags: | @UR029-1
|  | Given | no licence is registered for the current machine
|  | When | I start AssistDent
|  | And | I enter details including an EU country name
|  | Then | I can register
|  | And | AssistDent starts
TakePaymentEU:
| Tags: | @UR029-2
|  | Given | I have register a trial licence for an EU user
|  | When | I access the purchase licence page
|  | And | choose the default price
|  | Then | the payment is in Euros
|  | And | UK VAT is not applied
ViewInvoiceEuro:
| Tags: | @UR029-3
|  | Given | I have paid for a subscription for an EU user
|  | When | I access my billing portal
|  | Then | I can view an invoice for each month
|  | And | the invoice contains the price paid
|  | And | the invoice contains the customer name and address
|  | And | the invoice contains a statement regarding reverse tax

## CUSTOMERWORKFLO(2)
CustomerWorkflowUK
| Tags: | @UR, @UR028, @DesktopOnly, @Admin
As a company We want to register UK users and accept payment for subscriptions to AssistDent So we can make money
CanUseInUK
| Tags: | @UR028-1
|  | Given | no licence is registered for the current machine
|  | When | I start AssistDent
|  | And | I enter details including country name United Kingdom
|  | Then | Registration is successful
|  | And | AssistDent starts
PaymentOptionsUK
| Tags: | @UR028-2
|  | Given | I have register a trial licence for a UK user
|  | When | I access the purchase licence page
|  | And | choose the default price
|  | Then | the payment is in GBP
|  | And | UK VAT is applied
PurchaseLicence
| Tags: | @UR028-3
|  | Given | I have register a trial licence
|  | And | I record the license expiry date
|  | And | I access the purchase licence page
|  | And | choose a purchase option
|  | When | I complete the payment in Stripe
|  | And | I restart the app
|  | Then | the license expiry date is extended as specified by the product
SubscriptionDetails
| Tags: | @UR028-4
|  | Given | I have paid for a subscription
|  | When | I access my billing portal
|  | Then | I can see my subscription details
ViewInvoiceUK:
| Tags: | @UR028-5
|  | Given | I have paid for a subscription for a UK user
|  | When | I access my billing portal
|  | Then | I can view an invoice for each month
|  | And | the invoice contains the price paid
|  | And | the invoice contains the customer name and address
|  | And | the invoice contains the VAT amount
|  | And | the invoice contains the MIL VAT number
CancelSubscription
| Tags: | @UR028-6
|  | Given | I have paid for a subscription
|  | When | I access my billing portal
|  | And | Cancel my subsription
|  | And | I restart the app
|  | Then | the license expiry date remains unchanged
|  | And | my billing page shows that I have no active subscriptions
|  | And | the subscription is cancelled on Stripe

## DELETECARIESROI
DeleteCariesROI
| Tags: | @UR, @UR006, @Analysis
As a dentist I want to delete caries regions of interest that I judge to be incorrect So that I can share the results with a patient without showing false positives
DeleteCariesROI
| Tags: | @UR006-1
|  | Given | I submit an image
|  | And | the image contains caries ROIs
|  | When | I select one or more caries ROIs
|  | And | delete them
|  | Then | the caries ROIs are removed from the image

## DESKTOPLICENS(1)
DesktopLicensing
| Tags: | @UR, @UR025, @Admin, @DesktopOnly
As a company We want to issue users with a licence key So that we can restrict users to free trials and paying customers
RevokeLicence
| Tags: | @UR025-1
|  | Given | a licence is registered for the current machine
|  | And | I note the licece number and user ID
|  | When | I delete the settings file
|  | And | I delete license on Cryptolens
|  | And | I delete user on Stripe
|  | And | I restart AssistDent
|  | Then | I am asked to register
|  | And | AssistDent does not start
CannotUseIn
| Tags: | @UR025-2
|  | Given | no licence is registered for the current machine
|  | When | I start AssistDent
|  | And | I enter details including country name "United States"
|  | Then | I cannot register
|  | And | AssistDent does not start
RegisterTrialLicence
| Tags: | @UR025-3
|  | Given | no licence is registered for the current machine
|  | When | I start AssistDent
|  | Then | I am required to complete user registration
|  | And | I complete the registration
|  | And | the application starts
|  | And | a customer is created on CryptoLens with the correct details
|  | And | a customer is created on Stripe with the correct details
DisplayValidLicence
| Tags: | @UR025-4
|  | Given | a licence is registered for the current machine
|  | And | that licence has not expired
|  | When | I start the desktop app
|  | Then | the desktop app starts
|  | And | the licence key can be seen in the application
|  | And | the licence expiry date can be seen in the application
LicenceExpired
| Tags: | @UR025-5
|  | Given | a licence is registered for the current machine
|  | And | I set the licensence expiry date to yesterday on Cryptolens
|  | When | I start the desktop app
|  | Then | I am directed to the store website to purchase a Licence
|  | And | the app does not start
LicenceExtended
| Tags: | @UR025-6
|  | Given | a expired licence is registered for the current machine
|  | And | I set the licensence expiry date to tomorrow on Cryptolens
|  | When | I start the desktop app
|  | Then | the app starts
|  | And | the tomorrow is set as the licence expiry date can be seen in the application
EnterNewLicenceKey
| Tags: | @UR025-7
|  | Given | a valid licence is registered for the current machine
|  | And | I create a new license in cryptolens for the user
|  | And | I set the expiry date and correct product ID
|  | And | I record the new license key
|  | And | I enter the new license key in the app
|  | When | I restart the application
|  | Then | the application starts
|  | And | the application uses the new licence key
ChangeTrialPeriod
| Tags: | @UR025-8
|  | Given | a licence is not generated for the current machine
|  | When | I configure the store app to generate a 1 day licence
|  | And | I register for a trial licence
|  | Then | the licence expiry date can be seen in the application
|  | And | the expiry date is in 1 day

## DESKTOPUSAGEMETRICS
DesktopUsageMetrics
| Tags: | @UR, @UR032, @DesktopOnly, @Analysis
As a company We want to record the number of times AssistDent is used and the date of last use So that we can monitor take up of the product during evaluations
RecordNumberOfUses
| Tags: | @UR032-01
|  | Given | I have registered
|  | And | used AssistDent n times
|  | When | I view the information associated with my licence
|  | Then | that information shows AssistDent was used n times
RecordNumberOfUses
| Tags: | @UR032-02
|  | Given | I have registered
|  | And | used AssistDent today
|  | When | I view the information associated with my licence
|  | Then | that information shows AssistDent was last used today

## DISPLAYCARIESDETECTION
DisplayCariesDetection
| Tags: | @UR, @UR002, @Analysis
As a dentist I want to see which teeth have been detected and analyzed for caries So that I can check for caries manually on undetected teeth
TeethIndicatorsAreDisplayed
| Tags: | @UR002-1
|  | When | I submit an image
|  | Then | the detected teeth are indicated on the image

## DISPLAYCARIESDETECT(1)
DisplayCariesDetection
| Tags: | @UR, @UR003, @Analysis
As a dentist I want to view caries regions of interest on a bitewing image So that I can see where early stage decay is ocurring
ROIIndicatorsAreDisplayed
| Tags: | @UR003-1
|  | When | I submit an image
|  | Then | the detected caries ROIs are indicated
ROIIndicatorsAreEasilyVisible
| Tags: | @UR003-2
|  | When | I submit an image
|  | Then | the detected caries ROIs are indicated
|  | And | the detected caries ROI indicators are easily visible

## DISPLAYIM(1)
DisplayImage
| Tags: | @UR, @UR001, @Analysis
As a dentist I want to view one or more bitewing images So that when caries detection is performed I can view the results in place over the analyzed images
ImageIsDisplayed
| Tags: | @UR001-1
|  | When | I submit an image
|  | Then | the image is displayed on screen
MultipleImagesDisplayed
| Tags: | @UR001-2, @WebOnly
|  | When | I submit 4 images
|  | Then | all 4 images are displayed on screen

## ENTERFEEDBACK
EnterFeedback
| Tags: | @UR, @UR021, @Admin, @WebOnly
As a user I want to enter feedback comments So that I can tell Manchester Imaging what I think of assistdent.net
UserCanEnterFeedback
| Tags: | @UR021-1
|  | Given | I have logged in with analysis permissions
|  | When | I access the feedback page
|  | Then | I can enter a comment
AllowedIPCanEnterFeedback
| Tags: | @UR021-2
|  | Given | I have accessed from an allowed IP address
|  | When | I access the feedback page
|  | Then | I can enter a comment

## EXPANDIMAGE
ExpandImage
| Tags: | @UR, @UR005, @Analysis, @WebOnly
As a dentist I want to expand the image view to view full screen So that I can see potentially affected areas in more detail And I can share the image with a patient
ExpandImageToFullScreen
| Tags: | @UR005-1
|  | Given | I submit an image
|  | When | I expand the image view
|  | Then | the image is displayed using as much of the screen as possible
ExitFullScreen
| Tags: | @UR005-2
|  | Given | I am viewing an image in full screen
|  | When | I collapse the image view
|  | Then | the image is displayed in preview/tile view

## FEATURET(1)
FeatureTour
| Tags: | @UR, @UR012, @Labelling, @WebOnly
As a dentist I want to take a tour of features So that I can learn how to use assistdent.net
TakeTourOfFeatures
| Tags: | @UR012-1
|  | When | I follow a link to the tour
|  | Then | a tour of features is started
|  | And | submission is introduced
|  | And | display is introduced
|  | And | magnify is introduced
|  | And | hide/display annotation is introduced
|  | And | delete is introduced
|  | And | flip is introduced
|  | And | add is introduced
|  | And | full screen is introduced

## FOLDERWATC(1)
FolderWatcher
| Tags: | @UR, @UR030, @DesktopOnly, @Admin
As a dentist I want AssistDent to analyze any image added to a specified folder So that caries detection is performed automically on new images
AutomaticAnalysis
| Tags: | @UR030-01
|  | When | an image is copied into the watched folder
|  | Then | AssistDent automatically starts and analyzes the image
StartupAutomatically
| Tags: | @UR030-02
|  | When | logged into Windows
|  | Then | the Folder Watcher is started
|  | When | an image is copied into the watched folder
|  | Then | AssistDent automatically starts and analyzes the image
DisableFolderWatcher
| Tags: | @UR030-03
|  | Given | the Folder Watcher is enabled
|  | When | I disable the Folder Watcher
|  | And | an image is copied into the watched folder
|  | Then | AssistDent does not automatically start
ConfigureFolderWatcher
| Tags: | @UR030-04
|  | Given | the Folder Watcher is configured to one location
|  | And | a second location exists
|  | When | the watched folder is configured to the second location
|  | And | an image is copied into the watched folder
|  | Then | AssistDent automatically starts and analyzes the image

## IMAGEEXP(1)
ImageExport
| Tags: | @UR, @UR031, @DesktopOnly, @Analysis
As a dentist I want AssistDent to export the image with annotated results So that I can share the image with the patient and/or import it into my dental imaging software
ExportOnDemand
| Tags: | @UR031-01
|  | Given | I have performed caries detection
|  | When | I export an image
|  | Then | an image file is created
|  | And | that image contains the source X-Ray and annoted caries ROIs
ExportAutomatically
| Tags: | @UR031-02
|  | Given | I have performed caries detection
|  | And | I have enabled automatic export
|  | When | I close AssistDent
|  | Then | an image file is created
|  | And | that image contains the source X-Ray and annoted caries ROIs
SnapshotImage
| Tags: | @UR031-03
|  | Given | I have performed caries detection
|  | When | I snapshot an image
|  | Then | the clipboard contains an image
|  | And | that image contains the source X-Ray and annoted caries ROIs
ExportAutomatically
| Tags: | @UR031-04
|  | Given | I have performed caries detection
|  | And | I have configured AssistDent to run a command after export
|  | When | I export an image
|  | Then | the export is created
|  | And | the post-export command is executed

## IMAG(1)
ImageID
| Tags: | @UR, @UR023, @Analysis, @WebOnly
As a user I want to be able to see the filename name next to the image And be able to edit that ID to whatever text I wish So that I can identify and label any of the images on the analysis page And confirm that I submitted the correct image for analysis
ShowImageIDAsFilename
| Tags: | @UR023-1
|  | Given | I have submitted an image for analysis
|  | When | the image is displayed
|  | Then | the ID of the image is shown
|  | And | set as the image filename
EditImageID
| Tags: | @UR023-2
|  | Given | I have submitted an image for analysis
|  | When | I hover over the ID of the image
|  | Then | A box appears around the ID
|  | And | I am able to edit the ID field and change to different value

## LEGALLABELLING
LegalLabelling
| Tags: | @UR, @UR015, @Labelling
As a regulator I want to view the legal and safety information regarding AssistDent So that AssistDent can be used legally and safely on patients
DisplayLegalInformation
| Tags: | @UR015-1
|  | When | I request the clinical use information
|  | Then | the information includes the UDI Device Identifier
|  | And | the information includes the CE mark
|  | And | the information includes the manufacturer details
|  | And | the information includes the CE representative details
|  | And | the information includes the date of manufacture
|  | And | the information includes the intended of use
|  | And | the information includes the indications for use
|  | And | the information includes the contraindications
|  | And | the information includes the warnings and precautions

## LOGDETECTIONS
LogDetections
| Tags: | @UR, @UR019, @Security, @WebOnly
As an admin I want to view size, dimensions, analysis time and source IP address of detections by username So that I can see how often and when detections are being performed under each account As an admin I want to view size, dimensions, analysis time and source IP address of detections by Allowed IP So that I can see how often and when detections are being performed from an Allowed IP range
CanAccessUserDetections
| Tags: | @UR019-1
|  | Given | I have logged in with admin permissions
|  | When | I access the Details page for an individual user
|  | Then | I can see a list of detections
|  | And | I can see the date and time of the detection
|  | And | I can see the size and dimensions of the analysed image
|  | And | I can see the IP address from which the detection was made
CanAccessAllowedIPDetections
| Tags: | @UR019-2
|  | Given | I have logged in with admin permissions
|  | When | I access the Details page for an individual Allowed IP
|  | Then | I can see a list of detections
|  | And | I can see the date and time of the detection
|  | And | I can see the size and dimensions of the analysed image
|  | And | I can see the IP adderss from which the detection was made

## MAGNIFYIMAGE
MagnifyImage
| Tags: | @UR, @UR004, @Analysis
As a dentist I want to zoom in on a selected area of the image So that I can see potentially affected areas in more detail
MagnifierOnImage
| Tags: | @UR004-1
|  | Given | I submit an image
|  | And | the image is displayed on screen
|  | When | I hold the left mouse button down
|  | Then | a portion of the image is displayed at a higher magnification
HideAnnotationOnMagnify
| Tags: | @UR004-2
|  | Given | I submit an image
|  | And | I activate the magnifier
|  | When | I move the mouse cursor over a caries ROI indicator
|  | Then | the indicator is hidden in the magnified view
|  | When | I move the mouse cursor over a tooth indicator
|  | Then | the indicator is hidden in the magnified view

## MANAGEUSERLOG(1)
ManageUserLogins
| Tags: | @UR, @UR018, @Security, @WebOnly
As an administrator I want to create, view and delete user accounts So that I can grant individuals access to analysis functionality
ListUsers
| Tags: | @UR018-1
|  | Given | I am logged in with admin permissions
|  | When | I request the Manage Users page
|  | Then | a list of users is displayed
AddUser
| Tags: | @UR018-2
|  | Given | I am logged in with admin permissions
|  | When | I create an User
|  | And | choose a valid Password of at least 6 characters
|  | And | choose a valid Password with at least one lowercase, capital and number
|  | Then | my User is included in the list of users
DeleteUser
| Tags: | @UR018-3
|  | Given | I am logged in with admin permissions
|  | And | a single User exists
|  | When | I delete the User for my computer
|  | Then | the User is no longer present in the list of users

## PROTECTPATIENTDATA
ProtectPatientData
| Tags: | @UR, @UR011, @Security, @WebOnly
As a dentist I want to anonymise DICOM images before I send them to Manchester Imaging So that I can keep my patients' identities private As a dentist I want to encrypt image data when sent to Manchester Imaging So that I can prevent my patients' image data being intercepted en route
AnonymiseDicomData
| Tags: | @UR011-1
|  | When | I submit a DICOM image file for analysis
|  | Then | the DICOM file when read on the server contains no patient identifying data
SecureAllTraffic
| Tags: | @UR011-2
|  | When | I load the Assistdent web page
|  | Then | the page is accessible only via https
|  | And | the connection is secure

## REPOSITIONCARIESROI
RepositionCariesROI
| Tags: | @UR, @UR009, @Analysis
As a dentist I want to be able to reposition and change arrow indicator directions so that I can indicate alternative regions that I believe to be carious
MoveCariesROI
| Tags: | @UR009-2
|  | Given | I have submitted an image
|  | And | the image contains at least one ROI
|  | When | I click-and-drag with the mouse button on an ROI
|  | Then | the ROI is moved to the the new location
FlipCariesROI
| Tags: | @UR009-2
|  | Given | I submit an image
|  | And | the image contains caries ROIs
|  | When | I select one or more caries ROIs
|  | And | flip its direction
|  | Then | the caries ROIs direction is flipped horizontally

## SEEWHATS(1)
SeeWhatsNew
| Tags: | @UR, @UR027, @DesktopOnly, @Analysis
As a user I want to be made aware of new features So that I know what's changed
SeeNewFeaturesOnStartup
| Tags: | @UR027-1
|  | Given | I have not dismissed the news for the version installed
|  | When | I start AssistDent
|  | Then | the new features are displayed
StopSeeingNewFeatures
| Tags: | @UR027-2
|  | Given | I have not dismissed the news for the version installed
|  | And | I start AssistDent
|  | And | the new features are displayed
|  | When | I dismiss/request the new features are no longer displayed
|  | And | I restart AssistDent
|  | Then | AssistDent restarts
|  | And | the new features are not dismissed
SeeNewFeaturesAfterStartup
| Tags: | @UR027-3
|  | Given | I have started AssistDent
|  | When | I access the new features
|  | Then | the new features are displayed

## SHELLINTEGRAT(1)
ShellIntegration
| Tags: | @UR, @UR033, @DesktopOnly, @Admin
As a dentist I want to open image files with AssistDent So that I can use AssistDent with using a dental imaging app
OpenWith
| Tags: | @UR033-01
|  | Given | I have an image file
|  | When | in Explorer I right click, choose open with and find AssistDent.exe
|  | Then | the image opens in AssistDent
OpenWithMenuOption
| Tags: | @UR033-02
|  | Given | I have an image file
|  | When | in Explorer I right click, choose Open With AssistDent menu option
|  | Then | the image opens in AssistDent

## SHOWPROGRESS
ShowProgress
| Tags: | @UR, @UR026, @Analysis
As a user I want confirmation that AssistDent is performing caries detection So that I can wait for the analysis to complete
ShowAnalysisProgress
| Tags: | @UR026-1
|  | When | I perform caries detection
|  | Then | the analysis status is reported at relevant stages

## SUBMITIMAGES
SubmitImages
| Tags: | @UR, @UR016, @Analysis, @WebOnly
As a dentist I want to submit one or more images for analysis So that I can request caries detection on my patients' images
SubmitSingleImage
| Tags: | @UR016-1
|  | Given | I am authorized with analysis permissions
|  | And | I navigate to the Analysis page
|  | When | I browse for a single bitewing image
|  | Then | the image is sent
|  | And | the image is displayed
|  | And | results are displayed
SubmitMultipleImages
| Tags: | @UR016-2
|  | Given | I am authorized with analysis permissions
|  | And | I navigate to the Analysis page
|  | When | I browse and select 2 bitewing images
|  | Then | the images are sent
|  | And | the images are displayed
|  | And | results for both images are displayed
SubmitTiffImage
| Tags: | @UR016-3
|  | Given | I am authorized with analysis permissions
|  | And | I navigate to the Analysis page
|  | When | I browse for a single tiff bitewing image
|  | Then | the image is sent
|  | And | the image is displayed
|  | And | results are displayed
SubmitDICOMImage
| Tags: | @UR016-4
|  | Given | I am authorized with analysis permissions
|  | And | I navigate to the Analysis page
|  | When | I browse for a single DICOM bitewing image
|  | Then | the image is sent
|  | And | the image is displayed
|  | And | results are displayed

## TOGGLEANNOTATIONVISIBILITY
ToggleAnnotationVisibility
| Tags: | @UR, @UR007, @Analysis
As a dentist I want to hide all annotation So that I can review the image grade and look for other conditions
HideImageAnnotation
| Tags: | @UR007-1
|  | Given | I submit an image
|  | And | the image contains caries ROIs
|  | When | I hide the annotation
|  | Then | the tooth and caries ROI indicators are hidden from view
RestoreImageAnnotation
| Tags: | @UR007-2
|  | Given | I submit an image
|  | And | the image contains caries ROIs
|  | And | I have hidden the annotation
|  | When | I restore the annotation
|  | Then | the tooth and caries ROI indicators are displayed

## UPGRADECLIENT
UpgradeClient
| Tags: | @UR, @UR024, @Admin, @DesktopOnly
As a user I want to know when updates are available to the desktop client So that I can keep the software up to date And use the best performing algorithms
CheckForUpdatesNoneAvailable
| Tags: | @UR024-1
|  | Given | updates are not available
|  | When | I start the desktop app
|  | Then | the app starts normally
CheckForUpdatesWhenAvailable
| Tags: | @UR024-2
|  | Given | updates are available
|  | When | I start the desktop app
|  | Then | the user is informed of the availability of an update
DownloadUpdate
| Tags: | @UR024-3
|  | Given | updates are available
|  | When | I start the desktop app
|  | Then | the app starts the user is directed to the updated installer

## USERNAMELOGIN
UsernameLogin
| Tags: | @UR, @UR017, @Security, @WebOnly
As a dentist I want to authenticate myself using a username and password So that I can perform analyses from any remote computer
LoginWithUsername
| Tags: | @UR017-1
|  | Given | I am logged out
|  | And | I have Analysis user role
|  | When | I request login
|  | And | enter my username and password
|  | Then | I can access the analysis page
LoginWithUsernameWithoutAnalysisRole
| Tags: | @UR017-2
|  | Given | I am logged out
|  | And | I do not have Analysis user role
|  | When | I request login
|  | And | enter my username and password
|  | Then | I cannot access the Analysis page

## VIEWANALYSISPROGRESS
ViewAnalysisProgress
| Tags: | @UR, @UR022, @Analysis, @WebOnly
As a user I want to view the progress of my image upload, and the subsequent analysis So that I can see the application is responding to my request
UserCanViewUploadProgress
| Tags: | @UR022-1
|  | Given | I have logged in with analysis permissions
|  | When | I submit an image for analysis
|  | Then | I can see a progress indicator to show file upload progress
|  | And | it indicates how much of the file has been uploaded
UserCanViewAnalysisProgress
| Tags: | @UR022-2
|  | Given | I have logged in with analysis permissions
|  | And | I have submitted an image for analysis
|  | When | the file upload completes
|  | And | the image is visible
|  | Then | I can see a progress indicator to show analysis progress

## VIEWFEEDBACK
ViewFeedback
| Tags: | @UR, @UR020, @Admin, @WebOnly
As Manchester Imaging We want to view feedback comments by user or allowed ip So that we can see how a user/group of users feel about assistdent.net
CanAccessUserFeedback
| Tags: | @UR020-1
|  | Given | I have logged in with admin permissions
|  | When | I access the Details page for an individual user
|  | Then | I can see a list of feedback comments
|  | And | I can see the date of each comment
CanAccessAllowedIPFeedback
| Tags: | @UR020-2
|  | Given | I have logged in with admin permissions
|  | When | I access the Details page for an individual Allowed IP
|  | Then | I can see a list of feedback comments
|  | And | I can see the date of each comment