## TOC
functional
AddROI
DeleteROI
DisplayImage
ImageExport
ImageMagnifer
ProgressDisplay
RepositionROI
ROIIndicator
SeeWhatsNew
ToggleAnnotation
ToothIndicator

## ADDROI
AddROI
| Tags: | @FR, @FR008, @UR008, @Analysis
AddCariesROICursorAppearanceChange
| Tags: | @FR008-1, @UR008-1
|  | Given | I have submitted an image
|  | When | click the Add ROI button
|  | Then | the Add ROI button changes appearance
|  | And | the Cursor changes appearance to a crosshair
AddCariesROICursorAppearanceReturn
| Tags: | @FR008-2, @UR008-1
|  | Given | I have selected the Add ROI Button
|  | When | I hit select the Add ROI button again or any other button
|  | Then | the Add ROI button changes appearance back to original
|  | And | the Cursor changes appearance back to default
AddCariesROI
| Tags: | @FR008-3, @UR008-1
|  | Given | I have selected the Add ROI Button
|  | When | I move the cursor to the image and click
|  | Then | a new ROI is display with the arrow point at the mouse click
|  | And | pointing away from the nearest tooth centre
|  | And | the Add ROI button changes appearance back to original
|  | And | the Cursor changes appearance back to default

## DELETEROI
DeleteROI
| Tags: | @FR, @FR006, @UR006, @Analysis
DeleteROI
| Tags: | @FR006-2, @UR006-1
|  | Given | an analyzed bitewing image which contains caries ROIs
|  | When | I select an ROI
|  | And | click the delete button corresponding to that image
|  | Then | the selected caries ROI will be deleted
DeleteROIs
| Tags: | @FR006-3, @UR006-1
|  | Given | an analyzed bitewing image which contains caries ROIs
|  | When | I select 2 or more ROIs
|  | And | click the delete button corresponding to that image
|  | Then | the selected caries ROIs will be deleted

## DISPLAYIMAGE
DisplayImage
| Tags: | @FR, @FR001, @UR001, @Analysis
The image shall be displayed on screen
DisplayImage
| Tags: | @FR001-1, @UR001-1
|  | Given | a bitewing image
|  | When | submitted for analysis
|  | Then | the bitewing image is displayed on screen

## IMAGEEXPORT
ImageExport
| Tags: | @FR, @FR034, @UR031, @Analysis, @DesktopOnly
ExportImageButton
| Tags: | @FR034-01, @UR031-01
|  | Given | I have performed caries detection
|  | When | I click the Export button
|  | Then | an image file is created
|  | And | that image contains the source X-Ray and annoted caries ROIs
ExportROIsAndTeeth
| Tags: | @FR034-02, @UR031-01
|  | Given | I have performed caries detection
|  | And | I configure the export to include teeth
|  | When | I click the Export button
|  | Then | an image file is created
|  | And | that image contains the source X-Ray and annoted caries ROIs, and teeth indicators
ExportROIsOnly
| Tags: | @FR034-03, @UR031-01
|  | Given | I have performed caries detection
|  | And | I configure the export to exclude teeth
|  | When | I click the Export button
|  | Then | an image file is created
|  | And | that image contains the source X-Ray and annoted caries ROIs but no teeth
ExportToNewLocation
| Tags: | @FR034-04, @UR031-01
|  | Given | I have performed caries detection
|  | And | I configure the export to be created in a new folder
|  | When | I click the Export button
|  | Then | an image file is created in the new location
|  | And | that image contains the source X-Ray and annoted caries ROIs
ExportHasUniqueName
| Tags: | @FR034-05, @UR031-01
|  | Given | I have performed caries detection
|  | When | I click the Export button
|  | Then | an image file is created
|  | And | that image file has a unique name based on the date and time
ExportCreatedOnExit
| Tags: | @FR034-06, @UR031-02
|  | Given | I have performed caries detection
|  | And | I have configured AssistDent to export the image on exit
|  | When | I close AssistDent
|  | Then | an image file is created
|  | And | that image contains the source X-Ray and annoted caries ROIs
SnapshotImage
| Tags: | @FR034-07, @UR031-03
|  | Given | I have performed caries detection
|  | When | I click the Snapshot button
|  | Then | an image file is copied to the clipboard
|  | And | that image contains the source X-Ray and annoted caries ROIs
RunCommandAfterExport
| Tags: | @FR034-08, @UR031-04
|  | Given | I have performed caries detection
|  | And | I have configured AssisDent to run a command after export
|  | When | I click the Export button
|  | Then | the export is created
|  | And | the configured command is executed
ExportToVistaSoft
| Tags: | @FR034-09, @UR031-04
|  | Given | I have performed caries detection
|  | And | I have configured AssisDent to send to VistaSoft after export
|  | When | I click the Export button
|  | Then | the export is created
|  | And | the image appears alongside the parent image in VistaSoft

## IMAGEMAGNIFER
ImageMagnifer
| Tags: | @FR, @FR004, @UR004, @Analysis
With the mouse button held down, the section of the image directly under the mouse pointer will be zoomed to 3x magnification
ShowMagnifer
| Tags: | @FR004-1, @UR004-1
|  | Given | a bitewing image displayed on screen
|  | When | I hover the mouse pointer over a location in the image
|  | And | I hold the mouse button down
|  | Then | the magnifier should be displayed
|  | And | the image location should be displayed at 3x magnification
MoveMagnifer
| Tags: | @FR004-2, @UR004-1
|  | Given | a bitewing image displayed on screen
|  | And | the magnifier is currently displayed
|  | When | I move the mouse pointer to another location in the image
|  | Then | the image location should be updated
HideAnnotationInMagnifer
| Tags: | @FR004-3, @UR004-2
|  | Given | a bitewing image displayed on screen
|  | And | the magnifier is currently displayed
|  | And | the magnifer is positioned over an annotation
|  | Then | no annotation should be displayed in the magnification

## PROGRESSDISPLAY
ProgressDisplay
| Tags: | @FR, @FR028, @UR026, @Analysis, @DesktopOnly
ShowAnalysisFeedback
| Tags: | @FR028-1, @UR026-1
|  | When | I submit an image for analysis
|  | Then | the app displays status updates
|  | And | the updates are clearly visible

## REPOSITIONROI
RepositionROI
| Tags: | @FR, @FR009, @UR009, @Analysis
MoveAndDragROI
| Tags: | @FR009-1, @UR009-1
|  | Given | an analyzed bitewing image which contains caries ROIs
|  | When | I click and drag an ROI
|  | Then | the ROI will move with the mouse
|  | And | the ROI selection status will remain unchanged
MoveROI
| Tags: | @FR009-2, @UR009-1
|  | Given | an analyzed bitewing image which contains caries ROIs
|  | When | I click and drag an ROI
|  | And | move the mouse
|  | And | release the mouse
|  | Then | the ROI will stay at its new location
|  | And | the ROI selection status will remain unchanged
FlipROI
| Tags: | @FR009-3, @UR009-2
|  | Given | an analyzed bitewing image which contains caries ROIs
|  | When | I select an ROI
|  | And | click the flip button corresponding to that image
|  | Then | the selected caries ROI will be flipped
|  | And | the ROI will remain selected
FlipROIs
| Tags: | @FR009-4, @UR009-2
|  | Given | an analyzed bitewing image which contains caries ROIs
|  | When | I select 2 or more ROIs
|  | And | click the flip button corresponding to that image
|  | Then | the selected caries ROIs will be flipped
|  | And | the ROIs will remain selected

## ROIINDICATOR
ROIIndicator
| Tags: | @FR, @FR003, @UR003, @Analysis
The location of caries regions of interest shall be displayed in position on the image
ROIIndicators
| Tags: | @FR003-1, @UR003-1
|  | Given | a bitewing image known to contain caries
|  | When | caries detection is complete
|  | Then | caries ROIs are indicated over the image
ROIArrows
| Tags: | @FR003-2, @UR003-1
|  | Given | a bitewing image known to contain caries
|  | When | caries detection is complete
|  | Then | caries ROIs are indicated by arrows
|  | And | the arrows do not obscure the tooth edge
ROIEasilySeen
| Tags: | @FR003-3, @UR003-1
|  | Given | a bitewing image known to contain caries
|  | When | caries detection is complete
|  | Then | caries ROI indicators are a consistent size
|  | And | caries ROI indicators are brightly coloured

## SEEWHATSNEW
SeeWhatsNew
| Tags: | @FR, @FR029, @UR027, @Analysis, @DesktopOnly
SeeWhatsNewOnStartup
| Tags: | @FR029-1, @UR027-1
|  | Given | I have not dismissed the news for the version installed
|  | When | I start AssistDent
|  | Then | the what's new form is displayed
SeeWhatsNewAllVersions
| Tags: | @FR029-2, @UR027-1
|  | Given | an upgrade of AssistDent from the previous version
|  | And | the new features were previously dismissed
|  | When | I start AssistDent
|  | Then | the what's new form is displayed
|  | And | the what's new form contains the new features
|  | And | the new features apply only to the current version
SeeWhatsNewAllVersions
| Tags: | @FR029-3, @UR027-1
|  | Given | a clean install of AssistDent with no settings files
|  | When | I start AssistDent
|  | Then | the what's new form is displayed
|  | And | the what's new form contains the new features
|  | And | the new features apply to the current version and every previous version
DimissAndDontShowAgain
| Tags: | @FR029-4, @UR027-2
|  | Given | I have not dismissed the news for the version installed
|  | And | I start AssistDent
|  | And | the what's new form is displayed
|  | When | I check the "Don't show me this again" checkbox
|  | And | restart AssistDent
|  | Then | AssistDent starts
|  | And | the what's new form is not displayed
SeeWhatsNewAfterStartup
| Tags: | @FR029-5, @UR027-3
|  | Given | I have started AssistDent
|  | When | I access the "See What' New" menu option
|  | Then | the what's new form is displayed

## TOGGLEANNOTATION
ToggleAnnotation
| Tags: | @FR, @FR007, @UR007, @Analysis
HideImageAnnotation
| Tags: | @FR007-2, @UR007-1
|  | Given | an analyzed bitewing image which contains caries ROIs
|  | When | I click the corresponding hide annotation button
|  | Then | the tooth indicators are hidden
|  | And | the caries ROI indicators are hidden
|  | And | the hide annotation button is replaced with a show annotation button
RestoreImageAnnotation
| Tags: | @FR007-3, @UR007-2
|  | Given | an analyzed bitewing image which contains caries ROIs
|  | And | the annotation is currently hidden
|  | When | I click the corresponding show annotation button
|  | Then | the tooth indicators are displayed
|  | And | the caries ROI indicators are displayed
|  | And | the show annotation button is replaced with a hide annotation button

## TOOTHINDICATOR
ToothIndicator
| Tags: | @FR, @FR002, @UR002, @Analysis
Detected teeth shall be displayed in position over the image
ToothIndicators
| Tags: | @FR002-1, @UR002-1
|  | Given | a bitewing image known to contain teeth
|  | When | caries detection is complete
|  | Then | the detected teeth are displayed on the image