## TOC
user
AddCariesROI
DeleteCariesROI
DisplayCariesDetection
DisplayCariesDetection
DisplayImage
ExpandImage
ImageID
MagnifyImage
RepositionCariesROI
SubmitImages
ToggleAnnotationVisibility
ViewAnalysisProgress

## ADDCARIESROI
AddCariesROI
| Tags: | @UR, @UR008, @Analysis
As a dentist I want to add caries regions of interest at a location I deem dubious So that I can record for future use and share the results with a patient
AddCariesROI
| Tags: | @UR008-1
|  | Given | I submit an image
|  | When | I choose a location at which to add a caries ROI to the image
|  | Then | the new caries ROI is show on the image

## DELETECARIESROI
DeleteCariesROI
| Tags: | @UR, @UR006, @Analysis
As a dentist I want to delete caries regions of interest that I judge to be incorrect So that I can share the results with a patient without showing false positives
DeleteCariesROI
| Tags: | @UR006-1
|  | Given | I submit an image
|  | And | the image contains caries ROIs
|  | When | I select one or more caries ROIs
|  | And | delete them
|  | Then | the caries ROIs are removed from the image

## DISPLAYCARIESDETECTION
DisplayCariesDetection
| Tags: | @UR, @UR002, @Analysis
As a dentist I want to see which teeth have been detected and analyzed for caries So that I can check for caries manually on undetected teeth
TeethIndicatorsAreDisplayed
| Tags: | @UR002-1
|  | When | I submit an image
|  | Then | the detected teeth are indicated on the image

## DISPLAYCARIESDETECT(1)
DisplayCariesDetection
| Tags: | @UR, @UR003, @Analysis
As a dentist I want to view caries regions of interest on a bitewing image So that I can see where early stage decay is occuring
ROIIndicatorsAreDisplayed
| Tags: | @UR003-1
|  | When | I submit an image
|  | Then | the deteted caries ROIs are indicated
ROIIndicatorsAreEasilyVisible
| Tags: | @UR003-2
|  | When | I submit an image
|  | Then | the deteted caries ROIs are indicated
|  | And | the detected caries ROI indicators are easily visible

## DISPLAYIMAGE
DisplayImage
| Tags: | @UR, @UR001, @Analysis
As a dentist I want to view one or more bitewing images So that when caries detection is performed I can view the results in place over the analyzed images
ImageIsDisplayed
| Tags: | @UR001-1
|  | When | I submit an image
|  | Then | the image is displayed on screen
MultipleImagesDisplayed
| Tags: | @UR001-2
|  | When | I submit 4 images
|  | Then | all 4 images are displayed on screen

## EXPANDIMAGE
ExpandImage
| Tags: | @UR, @UR005, @Analysis
As a dentist I want to expand the image view to view full screen So that I can see potentially affected areas in more detail And I can share the image with a patient
ExpandImageToFullScreen
| Tags: | @UR005-1
|  | Given | I submit an image
|  | When | I expand the image view
|  | Then | the image is displayed using as much of the screen as possible
ExitFullScreen
| Tags: | @UR005-2
|  | Given | I am viewing an image in full screen
|  | When | I collapse the image view
|  | Then | the image is displayed in preview/tile view

## IMAGEID
ImageID
| Tags: | @UR, @UR023, @Analysis
As a user I want to be able to see the filename name next to the image And be able to edit that ID to whatever text I wish So that I can identify and label any of the images on the analysis page And confirm that I submitted the correct image for analysis
ShowImageIDAsFilename
| Tags: | @UR023-1
|  | Given | I have submitted an image for analysis
|  | When | the image is displayed
|  | Then | the ID of the image is shown
|  | And | set as the image filename
EditImageID
| Tags: | @UR023-2
|  | Given | I have submitted an image for analysis
|  | When | I hover over the ID of the image
|  | Then | A box appears around the ID
|  | And | I am able to edit the ID field and change to different value

## MAGNIFYIMAGE
MagnifyImage
| Tags: | @UR, @UR004, @Analysis
As a dentist I want to zoom in on a selected area of the image So that I can see potentially affected areas in more detail
MagnifierOnImage
| Tags: | @UR004-1
|  | Given | I submit an image
|  | And | the image is displayed on screen
|  | When | I hold the left mouse button down
|  | Then | a portion of the image is displayed at a higher magnification
HideAnnotationOnMagnify
| Tags: | @UR004-2
|  | Given | I submit an image
|  | And | I activate the magnifier
|  | When | I move the mouse cursor over a caries ROI indicator
|  | Then | the indicator is hidden in the magnified view
|  | When | I move the mouse cursor over a tooth indicator
|  | Then | the indicator is hidden in the magnified view

## REPOSITIONCARIESROI
RepositionCariesROI
| Tags: | @UR, @UR009, @Analysis
As a dentist I want to be able to reposition and change arrow indicator directions so that I can indicate alternative regions that I believe to be carious
MoveCariesROI
| Tags: | @UR009-2
|  | Given | I have submitted an image
|  | And | the image contains at least one ROI
|  | When | I click-and-drag with the mouse button on an ROI
|  | Then | the ROI is moved to the the new location
FlipCariesROI
| Tags: | @UR009-2
|  | Given | I submit an image
|  | And | the image contains caries ROIs
|  | When | I select one or more caries ROIs
|  | And | flip its direction
|  | Then | the caries ROIs direction is flipped horizontally

## SUBMITIMAGES
SubmitImages
| Tags: | @UR, @UR016, @Analysis
As a dentist I want to submit one or more images for analysis So that I can request caries detection on my patients' images
SubmitSingleImage
| Tags: | @UR016-1
|  | Given | I am authorized with analysis permissions
|  | And | I navigate to the Analysis page
|  | When | I browse for a single bitewing image
|  | Then | the image is sent
|  | And | the image is displayed
|  | And | results are displayed
SubmitMultipleImages
| Tags: | @UR016-2
|  | Given | I am authorized with analysis permissions
|  | And | I navigate to the Analysis page
|  | When | I browse and select 2 bitewing images
|  | Then | the images are sent
|  | And | the images are displayed
|  | And | results for both images are displayed
SubmitTiffImage
| Tags: | @UR016-3
|  | Given | I am authorized with analysis permissions
|  | And | I navigate to the Analysis page
|  | When | I browse for a single tiff bitewing image
|  | Then | the image is sent
|  | And | the image is displayed
|  | And | results are displayed
SubmitDICOMImage
| Tags: | @UR016-4
|  | Given | I am authorized with analysis permissions
|  | And | I navigate to the Analysis page
|  | When | I browse for a single DICOM bitewing image
|  | Then | the image is sent
|  | And | the image is displayed
|  | And | results are displayed

## TOGGLEANNOTATIONVISIBILITY
ToggleAnnotationVisibility
| Tags: | @UR, @UR007, @Analysis
As a dentist I want to hide all annotation So that I can review the image grade and look for other conditions
HideImageAnnotation
| Tags: | @UR007-1
|  | Given | I submit an image
|  | And | the image contains caries ROIs
|  | When | I hide the annotation
|  | Then | the tooth and caries ROI indicators are hidden from view
RestoreImageAnnotation
| Tags: | @UR007-2
|  | Given | I submit an image
|  | And | the image contains caries ROIs
|  | And | I have hidden the annotation
|  | When | I restore the annotation
|  | Then | the tooth and caries ROI indicators are displayed

## VIEWANALYSISPROGRESS
ViewAnalysisProgress
| Tags: | @UR, @UR022, @Analysis
As a user I want to view the progress of my image upload, and the subsequent analysis So that I can see the application is responding to my request
UserCanViewUploadProgress
| Tags: | @UR022-1
|  | Given | I have logged in with analysis permissions
|  | When | I submit an image for analysis
|  | Then | I can see a progress indicator to show file upload progress
|  | And | it indicates how much of the file has been uploaded
UserCanViewAnalysisProgress
| Tags: | @UR022-2
|  | Given | I have logged in with analysis permissions
|  | And | I have submitted an image for analysis
|  | When | the file upload completes
|  | And | the image is visible
|  | Then | I can see a progress indicator to show analysis progress