## TOC
system
functional
AddROI
AnalysisProgressIndicator
DeleteROI
DisplayImage
DisplayTextReport
DragAndDropExport
DragAndDropImport
EnableDentine
ImageExport
ImageID
ImageMagnifer
ProgressDisplay
RadiographQualityGrading
RepositionROI
ROIIndicator
TextReportExport
ToggleAnnotation
ToothIndicator

## ADDROI
AddROI
| Tags: | @FR, @FR008, @UR008, @Analysis
AddCariesROICursorAppearanceChange
| Tags: | @FR008-1, @UR008-1
|  | Given | I have submitted an image
|  | When | click the Add ROI button
|  | Then | the Add ROI button changes appearance
|  | And | the Cursor changes appearance to a cross hair
AddCariesROICursorAppearanceReturn
| Tags: | @FR008-2, @UR008-1
|  | Given | I have selected the Add ROI Button
|  | When | I hit select the Add ROI button again or any other button
|  | Then | the Add ROI button changes appearance back to original
|  | And | the Cursor changes appearance back to default
AddCariesROI
| Tags: | @FR008-3, @UR008-1
|  | Given | I have selected the Add ROI Button
|  | When | I move the cursor to the image and click
|  | Then | a new ROI is display with the arrow point at the mouse click
|  | And | pointing away from the nearest tooth centre
|  | And | the grade of the ROI is unknown
|  | And | the Add ROI button changes appearance back to original
|  | And | the Cursor changes appearance back to default
|  | And | the new caries ROI is listed in the text report
|  | And | the new caries ROI is listed as of unknown grade in the text report

## ANALYSISPROGRESSINDICATOR
AnalysisProgressIndicator
| Tags: | @FR, @FR024, @UR022, @Analysis
ShowAnalysisProgress
| Tags: | @FR024-2, @UR022-2
|  | Given | I have submitted an image for analysis
|  | When | the image is displayed
|  | Then | the status of the analysis is displayed over the image
|  | When | the analysis progresses
|  | Then | the status is updated to reflect the stage of the analysis

## DELETEROI
DeleteROI
| Tags: | @FR, @FR006, @UR006, @Analysis
DeleteROI
| Tags: | @FR006-2, @UR006-1
|  | Given | an analysed bitewing image which contains caries ROIs
|  | When | I select an ROI
|  | And | click the delete button corresponding to that image
|  | Then | the selected caries ROI will be deleted from the image
|  | And | the selected caries will be deleted from the text report
DeleteROIs
| Tags: | @FR006-3, @UR006-1
|  | Given | an analysed bitewing image which contains caries ROIs
|  | When | I select 2 or more ROIs
|  | And | click the delete button corresponding to that image
|  | Then | the selected caries ROIs will be deleted from the image
|  | And | the selected caries will be deleted from the text report

## DISPLAYIMAGE
DisplayImage
| Tags: | @FR, @FR001, @UR001, @Analysis
The image shall be displayed on screen
DisplayImage
| Tags: | @FR001-1, @UR001-1
|  | Given | a bitewing image
|  | When | submitted for analysis
|  | Then | the bitewing image is displayed on screen
DisplayDicomImage
| Tags: | @FR001-3, @UR001-1
|  | Given | a bitewing dicom image
|  | When | submitted for analysis
|  | Then | the bitewing image is displayed on screen
DisplayTiffImage
| Tags: | @FR001-4, @UR001-1
|  | Given | a bitewing tiff image
|  | When | submitted for analysis
|  | Then | the bitewing image is displayed on screen
DisplayJpegImage
| Tags: | @FR001-5, @UR001-1
|  | Given | a bitewing jpeg image
|  | When | submitted for analysis
|  | Then | the bitewing image is displayed on screen
DisplayPngImage
| Tags: | @FR001-6, @UR001-1
|  | Given | a bitewing png image
|  | When | submitted for analysis
|  | Then | the bitewing image is displayed on screen

## DISPLAYTEXTREPORT
DisplayTextReport
| Tags: | @FR, @FR040, @UR038, @Analysis
The text analyis report shall be displayed on screen and agree with the image display
DisplayTextReport
| Tags: | @FR040-1, @UR038-01
|  | Given | a bitewing image
|  | When | submitted for analysis
|  | Then | the bitewing analysis is displayed on the image
|  | And | the text version of the analysis is displayed on screen
|  | And | the text report agrees with the image display
DisplayTextReportEnamelOnly
| Tags: | @FR040-02, @UR038-01
|  | Given | I have performed caries detection on an image containing enamel and dentine caries
|  | And | include dentine caries is unchecked
|  | When | submitted for analysis
|  | Then | the bitewing analysis is displayed on the image
|  | And | the text version of the analysis is displayed on screen
|  | And | the text report contains enamel caries
|  | And | the text report does not contain dentine caries
DisplayTextReportEnamelAndDentine
| Tags: | @FR040-03, @UR038-01
|  | Given | I have performed caries detection on an image containing enamel and dentine caries
|  | And | include dentine caries is checked
|  | When | submitted for analysis
|  | Then | the bitewing analysis is displayed on the image
|  | And | the text version of the analysis is displayed on screen
|  | And | the text report contains enamel caries
|  | And | the text report contains dentine caries
TextReportCopyable
| Tags: | @FR040-04, @UR038-02
|  | Given | a bitewing image
|  | When | submitted for analysis
|  | Then | the bitewing analysis is displayed on the image
|  | And | the text version of the analysis is displayed on screen
|  | When | I highlight the text in the report
|  | And | I do CTRL+C
|  | And | I do CTRL+P into a text editor
|  | Then | the report text is copied to the text editor via the clipboard
ClinicalReportVisible
| Tags: | @FR040-05, @UR038-01
|  | Given | a bitewing image
|  | When | submitted for analysis
|  | Then | the bitewing analysis is displayed on the image
|  | And | the clinical report is displayed on screen
|  | And | the clinical report panel can be dragged and docked to the left edge
|  | And | the clinical report panel can be displayed as a window
|  | And | the clinical report cannot be closed

## DRAGANDDROPEXPORT
DragAndDropExport
| Tags: | @FR, @FR038, @UR036, @Analysis, @DesktopOnly
ShowDragCursor
| Tags: | @FR037-01, @UR036-01
|  | Given | an image has been analysed in AssistDent
|  | When | I click and drag the right-mouse button
|  | Then | the cursor is updated to show I am dragging something
DraggedExportFilename
| Tags: | @FR037-02, @UR036-01
|  | Given | an image has been analysed in AssistDent
|  | When | I click and drag the right-mouse button
|  | And | drop the image onto an app which supports drag and drop operations
|  | Then | a file appears in the app
|  | And | the filename incorporates the filename of the original image file
CanDropIntoOtherAppEnamelOnly
| Tags: | @FR037-03, @UR036-01
|  | Given | an image has been analysed in AssistDent which contains enamel and dentine caries
|  | And | include dentine caries is unchecked
|  | When | I click and drag the right-mouse button
|  | And | drop the image onto an app which supports drag and drop operations
|  | Then | a file appears in the app
|  | And | the file contains the x-ray image
|  | And | the file contains the enamel caries annotations and NOT the dentine annotations
CanDropIntoOtherAppEnamelAndDentine
| Tags: | @FR037-04, @UR036-01
|  | Given | an image has been analysed in AssistDent which contains enamel and dentine caries
|  | And | include dentine caries is checked
|  | When | I click and drag the right-mouse button
|  | And | drop the image onto an app which supports drag and drop operations
|  | Then | a file appears in the app
|  | And | the file contains the x-ray image
|  | And | the file contains the enamel caries annotations and the dentine annotations

## DRAGANDDROPIMPORT
DragAndDropImport
| Tags: | @FR, @FR037, @UR035, @Analysis, @DesktopOnly
AcceptDraggedImage
| Tags: | @FR037-01, @UR035-01
|  | Given | AssistDent is started
|  | When | I drag an image from another application
|  | And | drop the image into AssistDent
|  | Then | the image is loaded and analysed
HandleNonImage
| Tags: | @FR037-02, @UR035-01
|  | Given | AssistDent is started
|  | When | I drag a non-image file from another application
|  | And | drop the file into AssistDent
|  | Then | nothing happens
IgnoreMultipleImages
| Tags: | @FR037-03, @UR035-01
|  | Given | AssistDent is started
|  | When | I drag 2 or more images from another application
|  | And | drop the files into AssistDent
|  | Then | nothing happens
IgnoreWhenProcessing
| Tags: | @FR037-04, @UR035-01
|  | Given | AssistDent is in the middle analysing an image
|  | When | I drag an image from another application
|  | Then | AssistDent does not allow me to drop it in for analysis
AcceptAfterAnalysisCompleted
| Tags: | @FR037-05, @UR035-01
|  | Given | Analysis is completed
|  | When | I drag an image from another application
|  | Then | the image is loaded and analysed

## ENABLEDENTINE
EnableDentine
| Tags: | @FR, @FR039, @UR037, @Analysis, @DesktopOnly
ChangeDentineDisplay
| Tags: | @FR039-01, @UR037-01
|  | Given | I open the settings window
|  | And | change the option to include dentine caries
|  | When | I restart AssistDent
|  | And | open the settings window
|  | Then | my preference has been remembered
ShowDentineCaries
| Tags: | @FR039-02, @UR037-01
|  | Given | I have enabled the display of caries using the settings window
|  | When | I open an image containing both enamel and dentine caries
|  | Then | the dentine caries are indicated in the GUI
|  | And | the enamel caries are indicated in the GUI
|  | And | the dentine caries are indicated in the text report
|  | And | the enamel caries are indicated in the text report
HideDentineCaries
| Tags: | @FR039-03, @UR037-02
|  | Given | I have disabled the display of caries using the settings window
|  | When | I open an image containing both enamel and dentine caries
|  | Then | the dentine caries are not indicated in the GUI
|  | And | the enamel caries are indicated in the GUI
|  | And | the dentine caries are not indicated in the text report
|  | And | the enamel caries are indicated in the text report
AllowAssignmentOfDentineGrades
| Tags: | @FR039-04, @UR037-01
|  | Given | I have enabled the display of dentine caries using the settings window
|  | And | I open an image containing both enamel and dentine caries
|  | And | dentine caries are indicated
|  | When | I left-click on a dentine caries ROI
|  | Then | the grades menu is displayed
|  | And | the grade can be set to D
DenyAssignmentOfDentineGrades
| Tags: | @FR039-05, @UR037-02
|  | Given | I have disabled the display of dentine caries using the settings window
|  | And | I open an image containing both enamel and dentine caries
|  | When | I left-click on a caries ROI
|  | Then | the grades menu is displayed
|  | And | the grade cannot be set to D
DynamicAllowAssignmentOfDentineGrades
| Tags: | @FR039-06, @UR037-02
|  | Given | I have disabled the display of caries using the settings window
|  | When | I open an image containing both enamel and dentine caries
|  | When | I left-click on a caries ROI
|  | Then | the grades menu is displayed
|  | And | the grade cannot be set to D
|  | When | I enable dentine caries using the settings window
|  | When | I left-click on a caries ROI
|  | Then | the grades menu is displayed
|  | And | the grade can be set to D
DynamicDenyAssignmentOfDentineGrades
| Tags: | @FR039-07, @UR037-02
|  | Given | I have enabled the display of caries using the settings window
|  | When | I open an image containing both enamel and dentine caries
|  | When | I left-click on a caries ROI
|  | Then | the grades menu is displayed
|  | And | the grade can be set to D
|  | When | I disable dentine caries using the settings window
|  | When | I left-click on a caries ROI
|  | Then | the grades menu is displayed
|  | And | the grade cannot be set to D
DynamicShowDentineCariesUpdateGUI
| Tags: | @FR039-08, @UR037-03
|  | Given | I have disabled the display of caries using the settings window
|  | When | I open an image containing both enamel and dentine caries
|  | Then | the dentine caries are not indicated in the GUI
|  | And | the enamel caries are indicated in the GUI
|  | When | I enable dentine caries using the settings window
|  | Then | the dentine caries are indicated in the GUI
|  | And | the enamel caries are indicated in the GUI
DynamicHideDentineCariesUpdateGUI
| Tags: | @FR039-09, @UR037-04
|  | Given | I have enabled the display of caries using the settings window
|  | When | I open an image containing both enamel and dentine caries
|  | Then | the dentine caries are indicated in the GUI
|  | And | the enamel caries are indicated in the GUI
|  | When | I disable dentine caries using the settings window
|  | Then | the dentine caries are not indicated in the GUI
|  | And | the enamel caries are indicated in the GUI
DynamicShowDentineCariesUpdateTextReport
| Tags: | @FR039-10, @UR037-03
|  | Given | I have disabled the display of caries using the settings window
|  | When | I open an image containing both enamel and dentine caries
|  | Then | the dentine caries are indicated in the text report
|  | And | the enamel caries are indicated in the text report
|  | When | I enable dentine caries using the settings window
|  | Then | the dentine caries are indicated in the text report
|  | And | the enamel caries are indicated in the text report
DynamicHideDentineCariesUpdateTextReport
| Tags: | @FR039-11, @UR037-04
|  | Given | I have enabled the display of caries using the settings window
|  | When | I open an image containing both enamel and dentine caries
|  | Then | the dentine caries are indicated in the text report
|  | And | the enamel caries are indicated in the text report
|  | When | I disable dentine caries using the settings window
|  | Then | the dentine caries are not indicated in the text report
|  | And | the enamel caries are indicated in the text report

## IMAGEEXPORT
ImageExport
| Tags: | @FR, @FR034, @UR031, @Analysis, @DesktopOnly
ExportImageButton
| Tags: | @FR034-01, @UR031-01
|  | Given | I have performed caries detection
|  | When | I click the Export button
|  | Then | an image file is created
|  | And | that image contains the source X-Ray and annotated caries ROIs
ExportROIsEnamelOnly
| Tags: | @FR034-02, @UR031-01
|  | Given | I have performed caries detection on an image containing dentine caries
|  | And | include dentine caries is unchecked
|  | When | I click the Export button
|  | Then | an image file is created
|  | And | that image contains the source X-Ray and annotated enamel caries ROIs
|  | And | dentine caries ROI are NOT displayed
ExportROIsEnamelAndDentine
| Tags: | @FR034-03, @UR031-01
|  | Given | I have performed caries detection on an image containing dentine caries
|  | And | include dentine caries is checked
|  | When | I click the Export button
|  | Then | an image file is created
|  | And | that image contains the source X-Ray and annotated enamel caries ROIs
|  | And | dentine caries ROI are displayed
ExportToNewLocation
| Tags: | @FR034-04, @UR031-01
|  | Given | I have performed caries detection
|  | And | I configure the export to be created in a new folder
|  | When | I click the Export button
|  | Then | an image file is created in the new location
|  | And | that image contains the source X-Ray and annotated caries ROIs
ExportHasUniqueName
| Tags: | @FR034-05, @UR031-01
|  | Given | I have performed caries detection
|  | When | I click the Export button
|  | Then | an image file is created
|  | And | that image file has a unique name based on the date and time
|  | And | that name of the original input image is included in the name
ExportCreatedOnExit
| Tags: | @FR034-06, @UR031-02
|  | Given | I have performed caries detection
|  | And | I have configured AssistDent to export the image on exit
|  | When | I close AssistDent
|  | Then | an image file is created
|  | And | that image contains the source X-Ray and annotated caries ROIs
SnapshotImage
| Tags: | @FR034-07, @UR031-03
|  | Given | I have performed caries detection
|  | When | I click the Snapshot button
|  | Then | an image file is copied to the clipboard
|  | And | that image contains the source X-Ray and annotated caries ROIs
RunCommandAfterExport
| Tags: | @FR034-08, @UR031-04
|  | Given | I have performed caries detection
|  | And | I have configured AssisDent to run a command after export
|  | When | I click the Export button
|  | Then | the export is created
|  | And | the configured command is executed
ExportToVistaSoft
| Tags: | @FR034-09, @UR031-04
|  | Given | I have performed caries detection
|  | And | I have configured AssisDent to send to VistaSoft after export
|  | When | I click the Export button
|  | Then | the export is created
|  | And | the image appears alongside the parent image in VistaSoft
ExportAs
| Tags: | @FR034-10, @UR031-01
|  | Given | I have performed caries detection
|  | When | I click the ExportAs option
|  | Then | the file save dialog box opens
|  | And | the export image file is created
|  | And | the export image file is in the specified folder with the specified filename
ExportAsCancel
| Tags: | @FR034-11, @UR031-01
|  | Given | I have performed caries detection
|  | When | I click the ExportAs option
|  | And | I choose the cancel button
|  | Then | no export image file is created
ExportNoExam
| Tags: | @FR034-12, @UR031-01
|  | Given | I started AssistDent but not loaded an image for analysis
|  | When | I click the Export button
|  | Then | nothing happens
ExportAsNoExam
| Tags: | @FR034-13, @UR031-01
|  | Given | I started AssistDent but not loaded an image for analysis
|  | When | I click the ExportAs option
|  | Then | no file save dialog box opens
|  | And | nothing happens

## IMAGEID
ImageID
| Tags: | @FR, @FR025, @Analysis
ShowImageIDFilename
| Tags: | @FR025-1, @UR023-1
|  | Given | I have submitted an image for analysis
|  | Then | the image ID is shown
|  | And | the image ID displays the image filename

## IMAGEMAGNIFER
ImageMagnifer
| Tags: | @FR, @FR004, @UR004, @Analysis
With the mouse button held down, the section of the image directly under the mouse pointer will be zoomed to 3x magnification
ShowMagnifer
| Tags: | @FR004-1, @UR004-1
|  | Given | a bitewing image displayed on screen
|  | When | I hover the mouse pointer over a location in the image
|  | And | I hold the mouse button down
|  | Then | the magnifier should be displayed
|  | And | the image location should be displayed at 3x magnification
MoveMagnifer
| Tags: | @FR004-2, @UR004-1
|  | Given | a bitewing image displayed on screen
|  | And | the magnifier is currently displayed
|  | When | I move the mouse pointer to another location in the image
|  | Then | the image location should be updated
HideAnnotationInMagnifer
| Tags: | @FR004-3, @UR004-2
|  | Given | a bitewing image displayed on screen
|  | And | the magnifier is currently displayed
|  | And | the magnifier is positioned over an annotation
|  | Then | no annotation should be displayed in the magnification
MagnifierConsistentSize
| Tags: | @FR004-4, @UR004-1
|  | Given | a bitewing image file <image> with image width <width>
|  | When | caries detection is complete
|  | Then | GUI magnifier feature is a sensible and consistent size
Examples
|  |  | image | width
|  |  | TestData\images\testImages\variable_resolution\SC22_500_pixel_width.jpg | 500
|  |  | TestData\images\testImages\variable_resolution\SC22_1000_pixel_width.jpg | 1000
|  |  | TestData\images\testImages\variable_resolution\SC22_2000_pixel_width.jpg | 2000

## PROGRESSDISPLAY
ProgressDisplay
| Tags: | @FR, @FR028, @UR026, @Analysis, @DesktopOnly
ShowAnalysisFeedback
| Tags: | @FR028-1, @UR026-1
|  | When | I submit an image for analysis
|  | Then | the app displays status updates
|  | And | the updates are clearly visible

## RADIOGRAPHQUALITYGRADING
RadiographQualityGrading
| Tags: | @FR, @FR042, @UR040, @Analysis
ShowResolutionQualityMetric
| Tags: | @FR042-1, @UR040-1
|  | Given | a bitewing radiograph file <image> with resolution <width> and <height>
|  | When | AssistDent analyses the radiograph
|  | Then | the radiograph report displays the resolution <width> and <height>
|  | And | the resolution category <rating>
|  | And | the resolution pass/fail <status> based on the <width>
Examples
|  |  | image | width | height | rating | status
|  |  | TestData\images\testImages\variable_resolution\SC22_500_pixel_width.jpg | 500 | 389 | very low | fail
|  |  | TestData\images\testImages\variable_resolution\SC22_800_pixel_width.jpg | 800 | 622 | very low | fail
|  |  | TestData\images\testImages\variable_resolution\SC22_1000_pixel_width.jpg | 1000 | 778 | low | pass
|  |  | TestData\images\testImages\variable_resolution\SC22_1220_orig_width.jpg | 1220 | 950 | medium | pass
|  |  | TestData\images\testImages\variable_resolution\SC22_1500_pixel_width.jpg | 1500 | 1168 | medium | pass
|  |  | TestData\images\testImages\variable_resolution\SC22_1750_pixel_width.jpg | 1750 | 1362 | medium | pass
|  |  | TestData\images\testImages\variable_resolution\SC22_2000_pixel_width.jpg | 2000 | 1557 | high | pass
|  |  | TestData\images\testImages\variable_resolution\SC22_3000_pixel_width.jpg | 3000 | 2336 | very high | pass
ShowGrayScaleQualityMetric
| Tags: | @FR042-2, @UR040-2
|  | Given | a bitewing radiograph file <image> with gray scale flag <gray>
|  | When | AssistDent analyses the radiograph
|  | Then | the radiograph report displays the gray scale flag <gray>
|  | And | the gray scale pass/fail <status> based on the <gray> flag
Examples
|  |  | image | gray | status
|  |  | TestData\images\testImages\grayscale_vs_colour\SC22_exported.tif | FALSE | fail
|  |  | TestData\images\testImages\grayscale_vs_colour\SC22_orig.tif | TRUE | pass
ShowContrastQualityMetric
| Tags: | @FR042-3, @UR040-3
|  | Given | a bitewing radiograph file <image> with contrast metric <contrast>
|  | When | AssistDent analyses the radiograph
|  | Then | the radiograph report displays the contrast metric <contrast>
|  | And | the contrast category <rating>
|  | And | the contrast pass/fail <status> based on the <contrast>
Examples
|  |  | image | contrast | rating | status
|  |  | TestData\images\testImages\variable_contrast\SC22_0.2_contrast.jpg | 11 | very low | fail
|  |  | TestData\images\testImages\variable_contrast\SC22_0.5_contrast.jpg | 26 | very low | fail
|  |  | TestData\images\testImages\variable_contrast\SC22_0.75_contrast.jpg | 40 | low | pass
|  |  | TestData\images\testImages\variable_contrast\SC22_1.0_orig_contrast.jpg | 53 | medium | pass
|  |  | TestData\images\testImages\variable_contrast\SC22_1.2_contrast.jpg | 60 | high | pass
|  |  | TestData\images\testImages\variable_contrast\SC22_1.5_contrast.jpg | 69 | high | pass
|  |  | TestData\images\testImages\variable_contrast\SC22_2.5_contrast.jpg | 84 | very high | pass
|  |  | TestData\images\testImages\variable_contrast\SC22_3.0_contrast.jpg | 88 | very high | pass
ShowBlurQualityMetric
| Tags: | @FR042-4, @UR040-4
|  | Given | a bitewing radiograph file <image> with blur metric <blur>
|  | When | AssistDent analyses the radiograph
|  | Then | the radiograph report displays the blur metric <blur>
|  | And | the blur category <rating>
|  | And | the blur pass/fail <status> based on the <blur>
Examples
|  |  | image | blur | rating | status
|  |  | TestData\images\testImages\variable_blur\SC22_1.0_orig_blur.jpg | 51 | medium | pass
|  |  | TestData\images\testImages\variable_blur\SC22_1.5_blur.jpg | 72 | high | pass
|  |  | TestData\images\testImages\variable_blur\SC22_2.0_blur.jpg | 84 | very high | fail
|  |  | TestData\images\testImages\variable_blur\SC22_3.0_blur.jpg | 100 | very high | fail
|  |  | TestData\images\testImages\variable_blur\SC22_5.0_blur.jpg | 100 | very high | fail

## REPOSITIONROI
RepositionROI
| Tags: | @FR, @FR009, @UR009, @Analysis
MoveAndDragROI
| Tags: | @FR009-01, @UR009-1
|  | Given | an analysed bitewing image which contains caries ROIs
|  | When | I click and drag an ROI
|  | Then | the ROI and caries grade label will move with the mouse
|  | And | the ROI are deselected
|  | And | the text report is updated correctly based on the new caries location
MoveROIToNewToothDistal
| Tags: | @FR009-02, @UR009-1
|  | Given | an analysed bitewing image which contains distal edge caries ROIs
|  | When | I click and drag an ROI
|  | And | move the mouse to a new distal edge on a different tooth
|  | And | release the mouse
|  | Then | the ROI and grade label will stay at its new distal edge tooth location
|  | And | the ROI are deselected
|  | And | the ROI points away from the tooth centre to the distal edge
|  | And | the caries tooth number is updated in the text report based on the new distal edge tooth location
|  | And | the caries distal edge status is unchanged in the text report
MoveROIToNewToothMesial
| Tags: | @FR009-03, @UR009-1
|  | Given | an analysed bitewing image which contains mesial edge caries ROIs
|  | When | I click and drag an ROI
|  | And | move the mouse to a new mesial edge on a different tooth
|  | And | release the mouse
|  | Then | the ROI and grade label will stay at its new mesial edge tooth location
|  | And | the ROI are deselected
|  | And | the ROI points away from the tooth centre to the mesial edge
|  | And | the caries tooth number is updated in the text report based on the new mesial edge tooth location
|  | And | the caries mesial edge status is unchanged in the text report
MoveROIToNewToothDistalToMesial
| Tags: | @FR009-04, @UR009-1
|  | Given | an analysed bitewing image which contains distal edge caries ROIs
|  | When | I click and drag an ROI
|  | And | move the mouse to a new mesial edge on a different tooth
|  | And | release the mouse
|  | Then | the ROI and grade label will stay at its new mesial edge tooth location
|  | And | the ROI are deselected
|  | And | the ROI points away from the tooth centre to the mesial edge
|  | And | the caries tooth number is updated in the text report based on the new mesial edge tooth location
|  | And | the caries distal edge status is updated to mesial in the text report
MoveROIToNewToothMesialToDistal
| Tags: | @FR009-05, @UR009-1
|  | Given | an analysed bitewing image which contains mesial edge caries ROIs
|  | When | I click and drag an ROI
|  | And | move the mouse to a new distal edge on a different tooth
|  | And | release the mouse
|  | Then | the ROI and grade label will stay at its new distal edge tooth location
|  | And | the ROI are deselected
|  | And | the ROI points away from the tooth centre to the distal edge
|  | And | the caries tooth number is updated in the text report based on the new distal edge tooth location
|  | And | the caries mesial edge status is updated to distal in the text report
FlipROIUsingButton
| Tags: | @FR009-06, @UR009-2
|  | Given | an analysed bitewing image which contains caries ROIs
|  | When | I select an ROI
|  | And | click the flip button
|  | Then | the selected caries ROI will be flipped
|  | And | the ROI becomes deselected
|  | And | the caries tooth number is unchanged in the text report
|  | And | the caries distal/mesial edge status is updated in the text report based on the new direction
FlipROIsUsingButton
| Tags: | @FR009-07, @UR009-2
|  | Given | an analysed bitewing image which contains caries ROIs
|  | When | I select 2 or more ROIs
|  | And | click the flip button
|  | Then | the selected caries ROIs will be flipped
|  | And | the ROIs become deselected
|  | And | the caries tooth numbers are unchanged in the text report
|  | And | the caries distal/mesial edges are updated in the text report based on the new direction
FlipROIDistalToMesialUsingMouse
| Tags: | @FR009-08, @UR009-3
|  | Given | an analysed bitewing image which contains distal edge caries ROIs
|  | When | I click and drag an ROI arrow
|  | And | move the mouse to the mesial edge on the same tooth
|  | And | release the mouse
|  | Then | the ROI and grade label will stay at its new mesial edge tooth location
|  | And | the ROI are deselected
|  | And | the ROI points away from the tooth centre to the mesial edge
|  | And | the caries tooth number is unchanged in the text report
|  | And | the caries distal edge status is updated to mesial in the text report
FlipROIMesialToDistalUsingMouse
| Tags: | @FR009-09, @UR009-3
|  | Given | an analysed bitewing image which contains mesial edge caries ROIs
|  | When | I click and drag an ROI
|  | And | move the mouse to the distal edge on the same tooth
|  | And | release the mouse
|  | Then | the ROI and grade label will stay at its new distal edge tooth location
|  | And | the ROI are deselected
|  | And | the ROI points away from the tooth centre to the distal edge
|  | And | the caries tooth number is unchanged in the text report
|  | And | the caries mesial edge status is updated to distal in the text report
DoubleFlipROIUsingButton
| Tags: | @FR009-10, @UR009-2
|  | Given | an analysed bitewing image which contains caries ROIs
|  | When | I select an ROI
|  | And | click the flip button twice
|  | Then | the selected caries ROI will be flipped once
|  | And | the ROI becomes deselected
|  | And | the caries tooth number is unchanged in the text report
|  | And | the caries distal/mesial edge status is updated in the text report based on the new direction
DoubleFlipROIsUsingButton
| Tags: | @FR009-11, @UR009-2
|  | Given | an analysed bitewing image which contains caries ROIs
|  | When | I select 2 or more ROIs
|  | And | click the flip button twice
|  | Then | the selected caries ROI will be flipped once
|  | And | the ROIs become deselected
|  | And | the caries tooth numbers are unchanged in the text report
|  | And | the caries distal/mesial edges are updated in the text report based on the new direction

## ROIINDICATOR
ROIIndicator
| Tags: | @FR, @FR003, @UR003, @Analysis
The location of caries regions of interest shall be displayed in position on the image and in the text report.
ROIIndicators
| Tags: | @FR003-1, @UR003-1
|  | Given | a bitewing image known to contain caries
|  | When | caries detection is complete
|  | Then | caries ROIs are indicated over the image
|  | And | caries are listed in the text report
|  | And | caries are associated with the correct tooth in the text report
ROIArrows
| Tags: | @FR003-2, @UR003-1
|  | Given | a bitewing image known to contain caries
|  | When | caries detection is complete
|  | Then | caries ROIs are indicated by arrows
|  | And | the arrows do not obscure the tooth edge
|  | And | the arrow locations agree with the tooth labels listed in the text report
ROIEasilySeen
| Tags: | @FR003-3, @UR003-1
|  | Given | a bitewing image known to contain caries
|  | When | caries detection is complete
|  | Then | caries ROI indicators are a consistent size
|  | And | caries ROI indicators are brightly coloured
ROIConsistentSize
| Tags: | @FR003-4, @UR003-1
|  | Given | a bitewing image file <image> with image width <width>
|  | When | caries detection is complete
|  | Then | caries ROI GUI indicators are a sensible and consistent size
Examples
|  |  | image | width
|  |  | TestData\images\testImages\variable_resolution\SC22_500_pixel_width.jpg | 500
|  |  | TestData\images\testImages\variable_resolution\SC22_1000_pixel_width.jpg | 1000
|  |  | TestData\images\testImages\variable_resolution\SC22_2000_pixel_width.jpg | 2000
ROIGrades
| Tags: | @FR003-5, @UR003-2
|  | Given | a bitewing image known to contain caries of different grades
|  | When | caries detection is complete
|  | Then | caries ROI indicators for different grades are distinguishable
|  | And | caries grade is indicated in the text report.
EditCariesGrade
| Tags: | @FR003-6, @UR003-2, @DesktopOnly
|  | Given | caries ROI are displayed on the image and in the text report
|  | When | I click on the ROI label
|  | Then | a menu allowing me to change the grade appears
|  | And | I can choose a new grade
|  | And | the grade is updated with the new choice
|  | And | the grade is updated in the text report

## TEXTREPORTEXPORT
TextReportExport
| Tags: | @FR, @FR041, @UR039, @Analysis
ExportTextReportAndImageButton
| Tags: | @FR041-01, @UR039-01
|  | Given | I have performed caries detection
|  | When | I click the Export button
|  | Then | an analysis image file is created
|  | And | an analysis text report file is created
|  | And | the text report file agrees with the image file
|  | And | the text report contains a warning to review the exported image
ExportTextReportEnamelOnly
| Tags: | @FR041-02, @UR039-01
|  | Given | I have performed caries detection on an image containing enamel and dentine caries
|  | And | include dentine caries is unchecked
|  | When | I click the Export button
|  | Then | an text report file is created
|  | And | the text report contains enamel caries
|  | And | the text report does not contain dentine caries
|  | And | the text report contains a warning to review the exported image
ExportTextReportEnamelAndDentine
| Tags: | @FR041-03, @UR039-01
|  | Given | I have performed caries detection on an image containing dentine caries
|  | And | include dentine caries is checked
|  | When | I click the Export button
|  | Then | an text report file is created
|  | And | the text report contains enamel caries
|  | And | the text report contains dentine caries
|  | And | the text report contains a warning to review the exported image
ExportTextReportToNewLocation
| Tags: | @FR041-04, @UR039-01
|  | Given | I have performed caries detection
|  | And | I configure the export to be created in a new folder
|  | When | I click the Export button
|  | Then | an analysis image file is created in the new location
|  | And | an text report file is created in the new location
|  | And | the text report file agrees with the image file
|  | And | the text report contains a warning to review the exported image
ExportTextReportHasUniqueName
| Tags: | @FR041-05, @UR039-01
|  | Given | I have performed caries detection
|  | When | I click the Export button
|  | Then | an image file is created
|  | And | that image file has a unique name based on the date and time
|  | And | that name of the original input image is included in the name
|  | And | the text report file has the same name, but the file extension is ".txt"
|  | And | the text report file agrees with the image file
|  | And | the text report contains a warning to review the exported image
ExportTextReportCreatedOnExit
| Tags: | @FR041-06, @UR039-02
|  | Given | I have performed caries detection
|  | And | I have configured AssistDent to export the image on exit
|  | When | I close AssistDent
|  | Then | an image file is created
|  | And | a text report file is created
|  | And | the text report file agrees with the image file
|  | And | the text report contains a warning to review the exported image
ExportTextReportAs
| Tags: | @FR041-07, @UR039-01
|  | Given | I have performed caries detection
|  | When | I click the ExportAs option
|  | Then | the file save dialog opens
|  | And | the export image file is created
|  | And | the export image file is in the specified folder with the specified filename
|  | And | the text report file has the same name, but the file extension is ".txt"
|  | And | the text report file agrees with the image file
|  | And | the text report contains a warning to review the exported image
ExportTextReportAsCancel
| Tags: | @FR041-08, @UR039-01
|  | Given | I have performed caries detection
|  | When | I click the ExportAs option
|  | And | I choose the cancel button
|  | Then | no export image file is created
|  | And | no export text report file is created

## TOGGLEANNOTATION
ToggleAnnotation
| Tags: | @FR, @FR007, @UR007, @Analysis
HideImageAnnotation
| Tags: | @FR007-2, @UR007-1
|  | Given | an analysed bitewing image which contains caries ROIs
|  | When | I click the corresponding hide annotation button
|  | Then | the tooth indicators are hidden
|  | And | the caries ROI indicators are hidden
|  | And | the hide annotation button is replaced with a show annotation button
RestoreImageAnnotation
| Tags: | @FR007-3, @UR007-2
|  | Given | an analysed bitewing image which contains caries ROIs
|  | And | the annotation is currently hidden
|  | When | I click the corresponding show annotation button
|  | Then | the tooth indicators are displayed
|  | And | the caries ROI indicators are displayed
|  | And | the show annotation button is replaced with a hide annotation button

## TOOTHINDICATOR
ToothIndicator
| Tags: | @FR, @FR002, @UR002, @Analysis
Detected teeth shall be displayed in position over the image and the text report lists the detected teeth
ToothIndicators
| Tags: | @FR002-1, @UR002-1
|  | Given | a bitewing image known to contain adult teeth
|  | When | caries detection is complete
|  | Then | the detected teeth are displayed on the image
|  | And | identifiers for the teeth are shown
|  | And | the identifiers provide unique identification for each tooth
|  | And | the detected teeth are listed in the text report
ToothIndicators
| Tags: | @FR002-2, @UR002-1
|  | Given | a bitewing image file <image> with image width <width>
|  | When | caries detection is complete
|  | Then | the detected teeth are displayed on the image
|  | Then | caries ROI tooth labels are a sensible and consistent size
|  | And | the detected teeth are listed in the text report
Examples
|  |  | image | width
|  |  | TestData\images\testImages\variable_resolution\SC22_500_pixel_width.jpg | 500
|  |  | TestData\images\testImages\variable_resolution\SC22_1000_pixel_width.jpg | 1000
|  |  | TestData\images\testImages\variable_resolution\SC22_2000_pixel_width.jpg | 2000
EditToothQuadrant
| Tags: | @FR002-2, @UR002-2, @DesktopOnly
|  | Given | identifier for the teeth are displayed on the image
|  | When | I click on the tooth quadrant
|  | Then | a menu allowing me to change the quadrant appears
|  | And | I can choose a new quadrant
|  | And | the label is updated with the new quadrant value
|  | And | the tooth quadrant in the detected teeth list of the text report is also updated
EditToothNumber
| Tags: | @FR002-3, @UR002-3, @DesktopOnly
|  | Given | identifier for the teeth are displayed on the image
|  | When | I click on the tooth number
|  | Then | a menu allowing me to change the number appears
|  | And | I can choose a new number
|  | And | the label is updated with the new number value
|  | And | the tooth number in the detected teeth list of the text report is also updated