## TOC
functional
AddROI
AnalysisProgressIndicator
DeleteROI
DICOMImageFormat
DisplayImage
FullScreen
ImageID
ImageMagnifer
ImageSubmit
RepositionROI
ROIIndicator
TiffImageFormat
ToggleAnnotation
ToothIndicator

## ADDROI
AddROI
| Tags: | @FR, @FR008, @UR008, @Analysis
AddCariesROICursorAppearanceChange
| Tags: | @FR008-1, @UR008-1
|  | Given | I have submitted an image
|  | When | click the Add ROI button
|  | Then | the Add ROI button changes appearance
|  | And | the Cursor changes appearance to a crosshair
AddCariesROICursorAppearanceReturn
| Tags: | @FR008-2, @UR008-1
|  | Given | I have selected the Add ROI Button
|  | When | I hit select the Add ROI button again or any other button
|  | Then | the Add ROI button changes appearance back to original
|  | And | the Cursor changes appearance back to default
AddCariesROI
| Tags: | @FR008-3, @UR008-1
|  | Given | I have selected the Add ROI Button
|  | When | I move the cursor to the image and click
|  | Then | a new ROI is display with the arrow point at the mouse click
|  | And | pointing away from the nearest tooth centre
|  | And | the Add ROI button changes appearance back to original
|  | And | the Cursor changes appearance back to default

## ANALYSISPROGRESSINDICATOR
AnalysisProgressIndicator
| Tags: | @FR, @FR024, @UR022, @Analysis, @WebOnly
ShowFileUploadProgress
| Tags: | @FR024-1, @UR022-1, @WebOnly
|  | Given | I have logged in with analysis permissions
|  | When | I submit an image for analysis
|  | Then | I can see a progress bar
|  | And | the progress bar fills from left to right until completely full
|  | When | the progress bar is full
|  | Then | the image is displayed
|  | And | the progress bar is hidden
ShowAnalysisProgress
| Tags: | @FR024-2, @UR022-2
|  | Given | I have submitted an image for analysis
|  | When | the image is displayed
|  | Then | the status of the analysis is displayed over the image
|  | When | the analysis progresses
|  | Then | the status is updated to reflect the stage of the analysis

## DELETEROI
DeleteROI
| Tags: | @FR, @FR006, @UR006, @Analysis
DisplayDeleteButton
| Tags: | @FR006-1, @UR006-1, @WebOnly
|  | Given | 2 submitted bitewing images
|  | Then | the images shall be displayed in tile view
|  | And | a delete button shall be displayed alongside each image
DeleteROI
| Tags: | @FR006-2, @UR006-1
|  | Given | an analyzed bitewing image which contains caries ROIs
|  | When | I select an ROI
|  | And | click the delete button corresponding to that image
|  | Then | the selected caries ROI will be deleted
DeleteROIs
| Tags: | @FR006-3, @UR006-1
|  | Given | an analyzed bitewing image which contains caries ROIs
|  | When | I select 2 or more ROIs
|  | And | click the delete button corresponding to that image
|  | Then | the selected caries ROIs will be deleted

## DICOMIMAGEFORMAT
DICOMImageFormat
| Tags: | @FR, @FR017, @UR016, @Analysis, @WebOnly
CanAnalyseTiffImages
| Tags: | @FR017-1, @UR016-1, @UR016-4
|  | Given | I have logged in with analysis permissions
|  | When | I enter URL <base>/Analysis
|  | And | click the Select File area
|  | And | browse for a a bitewing image in DICOM format
|  | And | submit the image
|  | Then | the image is displayed correctly
|  | And | the results are displayed

## DISPLAYIMAGE
DisplayImage
| Tags: | @FR, @FR001, @UR001, @Analysis
The image shall be displayed on screen
DisplayImage
| Tags: | @FR001-1, @UR001-1
|  | Given | a bitewing image
|  | When | submitted for analysis
|  | Then | the bitewing image is displayed on screen
DisplayImages
| Tags: | @FR001-2, @UR001-2, @WebOnly
|  | Given | multiple bitewing images
|  | When | submitted for analysis
|  | Then | all bitewing images are displayed on screen

## FULLSCREEN
FullScreen
| Tags: | @FR, @FR005, @UR005, @Analysis, @WebOnly
The image shall be displayed expanded to fill as much of the screen as possible
AccessFullScreen
| Tags: | @FR005-1, @UR005-1
|  | Given | 2 submitted bitewing images
|  | Then | the images shall be displayed in tile view
|  | And | an expand button shall be displayed alongside each image
EnterFullScreen
| Tags: | @FR005-2, @UR005-1
|  | Given | 2 images displayed in tile view
|  | When | I press the the expand button corresponding to the first image
|  | Then | the first image is displayed to fill as much of the screen as possible
|  | And | the other image is hidden from view
AccessTileView
| Tags: | @FR005-3, @UR005-1
|  | Given | an image displayed full screen
|  | Then | an exit/collapse button shall be displayed alongside the image
ReturnToTileView
| Tags: | @FR005-4, @UR005-1
|  | Given | an image displayed full screen
|  | When | I press the exit/collapse button the image is displayed in tile view

## IMAGEID
ImageID
| Tags: | @FR, @FR025, @Analysis, @WebOnly
ShowImageIDFilename
| Tags: | @FR025-1, @UR023-1
|  | Given | I have submitted an image for analysis
|  | Then | the image ID is shown
|  | And | the image ID displays the image filename
ShowImageIDEditBox
| Tags: | @FR025-2, @UR023-2
|  | Given | I have submitted an image for analysis
|  | And | the image ID displays the image filename
|  | When | I hover over the ID
|  | Then | a box is shown indicating that the field can be edited

## IMAGEMAGNIFER
ImageMagnifer
| Tags: | @FR, @FR004, @UR004, @Analysis
With the mouse button held down, the section of the image directly under the mouse pointer will be zoomed to 3x magnification
ShowMagnifer
| Tags: | @FR004-1, @UR004-1
|  | Given | a bitewing image displayed on screen
|  | When | I hover the mouse pointer over a location in the image
|  | And | I hold the mouse button down
|  | Then | the magnifier should be displayed
|  | And | the image location should be displayed at 3x magnification
MoveMagnifer
| Tags: | @FR004-2, @UR004-1
|  | Given | a bitewing image displayed on screen
|  | And | the magnifier is currently displayed
|  | When | I move the mouse pointer to another location in the image
|  | Then | the image location should be updated
HideAnnotationInMagnifer
| Tags: | @FR004-3, @UR004-2
|  | Given | a bitewing image displayed on screen
|  | And | the magnifier is currently displayed
|  | And | the magnifer is positioned over an annotation
|  | Then | no annotation should be displayed in the magnification

## IMAGESUBMIT
ImageSubmit
| Tags: | @FR, @FR018, @UR016, @Analysis, @WebOnly
CanSubmitImageByBrowsing
| Tags: | @FR018-1, @UR016-1
|  | Given | I have logged in with analysis permissions
|  | When | I enter URL <base>/Analysis
|  | And | click the dashed submit area
|  | And | browse for a bitewing image
|  | And | submit the image
|  | Then | the image is displayed correctly
|  | And | the results are displayed
CanSubmitImageByDragAndDrop
| Tags: | @FR018-1, @UR016-2
|  | Given | I have logged in with analysis permissions
|  | When | I enter URL <base>/Analysis
|  | And | drag and drop a bitewing image from Explorer to the dashed submit area
|  | Then | the image is displayed correctly
|  | And | the results are displayed
CanSubmitImagesByBrowsing
| Tags: | @FR018-3, @UR016-1
|  | Given | I have logged in with analysis permissions
|  | When | I enter URL <base>/Analysis
|  | And | click the dashed submit area
|  | And | browse for 2 bitewing images
|  | And | submit the images
|  | Then | the images are displayed correctly
|  | And | the results are displayed
CanSubmitImagesByDragAndDrop
| Tags: | @FR018-4, @UR016-2
|  | Given | I have logged in with analysis permissions
|  | When | I enter URL <base>/Analysis
|  | And | drag and drop 2 bitewing images from Explorer to the dashed submit area
|  | Then | the images are displayed correctly
|  | And | the results are displayed

## REPOSITIONROI
RepositionROI
| Tags: | @FR, @FR009, @UR009, @Analysis
MoveAndDragROI
| Tags: | @FR009-1, @UR009-1
|  | Given | an analyzed bitewing image which contains caries ROIs
|  | When | I click and drag an ROI
|  | Then | the ROI will move with the mouse
|  | And | the ROI selection status will remain unchanged
MoveROI
| Tags: | @FR009-2, @UR009-1
|  | Given | an analyzed bitewing image which contains caries ROIs
|  | When | I click and drag an ROI
|  | And | move the mouse
|  | And | release the mouse
|  | Then | the ROI will stay at its new location
|  | And | the ROI selection status will remain unchanged
FlipROI
| Tags: | @FR009-3, @UR009-2
|  | Given | an analyzed bitewing image which contains caries ROIs
|  | When | I select an ROI
|  | And | click the flip button corresponding to that image
|  | Then | the selected caries ROI will be flipped
|  | And | the ROI will remain selected
FlipROIs
| Tags: | @FR009-4, @UR009-2
|  | Given | an analyzed bitewing image which contains caries ROIs
|  | When | I select 2 or more ROIs
|  | And | click the flip button corresponding to that image
|  | Then | the selected caries ROIs will be flipped
|  | And | the ROIs will remain selected

## ROIINDICATOR
ROIIndicator
| Tags: | @FR, @FR003, @UR003, @Analysis
The location of caries regions of interest shall be displayed in position on the image
ROIIndicators
| Tags: | @FR003-1, @UR003-1
|  | Given | a bitewing image known to contain caries
|  | When | caries detection is complete
|  | Then | caries ROIs are indicated over the image
ROIArrows
| Tags: | @FR003-2, @UR003-1
|  | Given | a bitewing image known to contain caries
|  | When | caries detection is complete
|  | Then | caries ROIs are indicated by arrows
|  | And | the arrows do not obscure the tooth edge
ROIEasilySeen
| Tags: | @FR003-3, @UR003-1
|  | Given | a bitewing image known to contain caries
|  | When | caries detection is complete
|  | Then | caries ROI indicators are a consistent size
|  | And | caries ROI indicators are brightly coloured
ROIConsistentSize
| Tags: | @FR003-4, @UR003-1, @WebOnly
|  | Given | 2 bitewing images of different resolution
|  | When | caries detection is complete
|  | Then | caries ROI indicators in both images are a consistent size

## TIFFIMAGEFORMAT
TiffImageFormat
| Tags: | @FR, @FR016, @UR016, @Analysis, @WebOnly
CanAnalyseTiffImages
| Tags: | @FR016-1, @UR016-1, @UR016-3
|  | Given | I have logged in with analysis permissions
|  | When | I enter URL <base>/Analysis
|  | And | click the Select File area
|  | And | browse for a a bitewing image in Tiff format
|  | And | submit the image
|  | Then | the image is displayed correctly
|  | And | the results are displayed

## TOGGLEANNOTATION
ToggleAnnotation
| Tags: | @FR, @FR007, @UR007, @Analysis
AccessAnnotationToggle
| Tags: | @FR007-1, @UR007-1, @WebOnly
|  | Given | 2 submitted bitewing images
|  | Then | the images shall be displayed in tile view
|  | And | a hide annotation button shall be displayed alongside each image
HideImageAnnotation
| Tags: | @FR007-2, @UR007-1
|  | Given | an analyzed bitewing image which contains caries ROIs
|  | When | I click the corresponding hide annotation button
|  | Then | the tooth indicators are hidden
|  | And | the caries ROI indicators are hidden
|  | And | the hide annotation button is replaced with a show annotation button
RestoreImageAnnotation
| Tags: | @FR007-3, @UR007-2
|  | Given | an analyzed bitewing image which contains caries ROIs
|  | And | the annotation is currently hidden
|  | When | I click the corresponding show annotation button
|  | Then | the tooth indicators are displayed
|  | And | the caries ROI indicators are displayed
|  | And | the show annotation button is replaced with a hide annotation button

## TOOTHINDICATOR
ToothIndicator
| Tags: | @FR, @FR002, @UR002, @Analysis
Detected teeth shall be displayed in position over the image
ToothIndicators
| Tags: | @FR002-1, @UR002-1
|  | Given | a bitewing image known to contain teeth
|  | When | caries detection is complete
|  | Then | the detected teeth are displayed on the image