## TOC
user
AddCariesROI
DeleteCariesROI
DesktopUsageMetrics
DisplayCariesDetection
DisplayImage
DisplayShortTextReport
DisplayTeethDetection
DragAndDropExport
DragAndDropImport
EnableDentine
ImageExport
ImageID
MagnifyImage
RadiographQualityGrading
RepositionCariesROI
ShowProgress
TextReportExport
ToggleAnnotationVisibility

## ADDCARIESROI
AddCariesROI
| Tags: | @UR, @UR008, @Analysis
As a dentist I want to add caries regions of interest at a location that is suspected to be carious So that I can record for future use and share the results with a patient
AddCariesROI
| Tags: | @UR008-1
|  | Given | I submit an image
|  | When | I choose a location at which to add a caries ROI to the image
|  | Then | the new caries ROI is show on the image
|  | And | the new caries ROI is labelled as of unknown grade in the image
|  | And | the new caries ROI is included in the text report
|  | And | the new caries ROI is labelled as of unknown grade in the text report

## DELETECARIESROI
DeleteCariesROI
| Tags: | @UR, @UR006, @Analysis
As a dentist I want to delete caries regions of interest that I judge to be incorrect So that I can share the results with a patient without showing false positives
DeleteCariesROI
| Tags: | @UR006-1
|  | Given | I submit an image
|  | And | the image contains caries ROIs
|  | When | I select one or more caries ROIs
|  | And | delete them
|  | Then | the caries ROIs are removed from the image
|  | And | the caries are removed from the text report

## DESKTOPUSAGEMETRICS
DesktopUsageMetrics
| Tags: | @UR, @UR032, @DesktopOnly, @Analysis
As a company We want to record the number of times AssistDent is used and the date of last use So that we can monitor take up of the product during evaluations
RecordNumberOfUses
| Tags: | @UR032-01
|  | Given | I have registered
|  | And | used AssistDent n times
|  | When | I view the information associated with my licence that incorporates my registered name
|  | Then | that information shows AssistDent was used n times
RecordNumberOfUses
| Tags: | @UR032-02
|  | Given | I have registered
|  | And | used AssistDent today
|  | When | I view the information associated with my licence that incorporates my registered name
|  | Then | that information shows AssistDent was last used today

## DISPLAYCARIESDETECTION
DisplayCariesDetection
| Tags: | @UR, @UR003, @Analysis
As a dentist I want to view caries regions of interest on a bitewing image and list the caries in the text report So that I can see where caries is occurring
ROIIndicatorsAreDisplayed
| Tags: | @UR003-1
|  | When | I submit an image
|  | Then | the detected caries ROIs are indicated in the image
|  | And | the detected caries are listed in the text report
|  | And | the detected caries are associated with the correct tooth label in the text report
ROIIndicatorsAreEasilyVisible
| Tags: | @UR003-2
|  | When | I submit an image
|  | Then | the detected caries ROIs are indicated in the image
|  | And | the detected caries ROI indicators are easily visible
|  | And | the detected caries ROI indicators for different grades are distinguishable
|  | And | the detected caries grades are listed in the text report

## DISPLAYIMAGE
DisplayImage
| Tags: | @UR, @UR001, @Analysis
As a dentist I want to view one or more bitewing images So that when caries detection is performed I can view the results in place over the analysed images
ImageIsDisplayed
| Tags: | @UR001-1
|  | When | I submit an image
|  | Then | the image is displayed on screen

## DISPLAYSHORTTEXTREPORT
DisplayShortTextReport
| Tags: | @UR, @UR038, @Analysis, @DesktopOnly
As a dentist I want a short text report displayed that agrees with the image analysis display So that when caries detection is performed I can easily copy the caries analysis into my clinical notes
TextReportDisplayed
| Tags: | @UR038-1
|  | When | I submit an image
|  | Then | the analysis result is displayed on the image
|  | And | a text report is displayed
|  | And | the text report agrees with the image analysis display
TextReportCopyable
| Tags: | @UR038-2
|  | When | I submit an image
|  | Then | the analysis result is displayed on the image
|  | And | a text report is displayed
|  | When | I highlight the text in the report
|  | Then | I can copy the text to a text editor.

## DISPLAYTEETHDETECTION
DisplayTeethDetection
| Tags: | @UR, @UR002, @Analysis
As a dentist I want to see which teeth have been detected and analysed for caries So that I can check for caries manually on undetected teeth
TeethIndicatorsAreDisplayed
| Tags: | @UR002-1
|  | When | I submit an image
|  | Then | the detected teeth are indicated on the image
|  | And | identifiers for the teeth are shown
EditToothQuadrant
| Tags: | @UR002-2, @DesktopOnly
|  | Given | teeth identifiers are show
|  | Then | I can edit the tooth quadrant
|  | And | the tooth quadrant in the text report is updated
EditToothNumber
| Tags: | @UR002-3, @DesktopOnly
|  | Given | teeth identifiers are show
|  | Then | I can edit the tooth number
|  | And | the tooth number in the text report is updated

## DRAGANDDROPEXPORT
DragAndDropExport
| Tags: | @UR, @UR036, @DesktopOnly, @Analysis
As a dentist I want to drag an image, including annotation, from AssistDent into my dental imaging system So that I can keep a record of the caries detection results
ExportByDragAndDrop
| Tags: | @UR036-01
|  | Given | I have analysed an image in AssistDent
|  | When | I drag the image out of AssistDent
|  | And | drop the image into another application that supports the drop operation
|  | Then | the image is loaded into the other application

## DRAGANDDROPIMPORT
DragAndDropImport
| Tags: | @UR, @UR035, @DesktopOnly, @Analysis
As a dentist I want to drag an image from my dental imaging system into AssistDent So that I can perform caries detection on my patients images
ImportByDragAndDrop
| Tags: | @UR035-01
|  | Given | AssistDent is running
|  | When | I drag an image from another application
|  | And | drop the image into AssistDent
|  | Then | AssistDent loads and analyses the dropped image

## ENABLEDENTINE
EnableDentine
| Tags: | @UR, @UR037, @DesktopOnly, @Analysis
As a dentist I want to enable and disable the display and incorporation of dentine caries in my analysis So that I can optionally concentrate on just enamel-only proximal caries
EnableDentine
| Tags: | @UR037-01
|  | Given | I have enabled dentine
|  | When | I open an image which contains enamel and dentine caries
|  | Then | both enamel and dentine caries are displayed
|  | And | I can change the grade of a caries ROI to D
|  | And | the text report contains enamel and dentine caries
DisableDentine
| Tags: | @UR037-02
|  | Given | I have disabled dentine
|  | When | I open an image which contains enamel and dentine caries
|  | Then | only enamel caries are displayed
|  | And | I am unable to change the grade to D
|  | And | the text report contains only enamel caries
EnableDentineUpdateCurrentAnalysis
| Tags: | @UR037-03
|  | Given | I have disabled dentine
|  | When | I open an image which contains enamel and dentine caries
|  | Then | only enamel caries are displayed
|  | And | the text report contains only enamel caries
|  | When | I enable dentine
|  | Then | the dentine caries appear in the analysis GUI
|  | And | the text report contains enamel and dentine caries
DisableDentineUpdateCurrentAnalysis
| Tags: | @UR037-04
|  | Given | I have enabled dentine
|  | When | I open an image which contains enamel and dentine caries
|  | Then | enamel and dentine caries are displayed
|  | And | the text report contains enamel and dentine caries
|  | When | I disable dentine
|  | Then | only enamel caries appear in the analysis GUI
|  | And | the text report contains only enamel caries

## IMAGEEXPORT
ImageExport
| Tags: | @UR, @UR031, @DesktopOnly, @Analysis
As a dentist I want AssistDent to export the image with annotated results So that I can share the image with the patient and/or import it into my dental imaging software
ExportOnDemand
| Tags: | @UR031-01
|  | Given | I have performed caries detection
|  | When | I export an image
|  | Then | an image file is created
|  | And | that image contains the source X-Ray and annotated caries ROIs
ExportAutomatically
| Tags: | @UR031-02
|  | Given | I have performed caries detection
|  | And | I have enabled automatic export
|  | When | I close AssistDent
|  | Then | an image file is created
|  | And | that image contains the source X-Ray and annotated caries ROIs
SnapshotImage
| Tags: | @UR031-03
|  | Given | I have performed caries detection
|  | When | I snapshot an image
|  | Then | the clipboard contains an image
|  | And | that image contains the source X-Ray and annotated caries ROIs
RunCommandAfterExport
| Tags: | @UR031-04
|  | Given | I have performed caries detection
|  | And | I have configured AssistDent to run a command after export
|  | When | I export an image
|  | Then | the export is created
|  | And | the post-export command is executed

## IMAGEID
ImageID
| Tags: | @UR, @UR023, @Analysis
As a user I want to be able to see the filename name next to the image And be able to edit that ID to whatever text I wish So that I can identify and label any of the images on the analysis page And confirm that I submitted the correct image for analysis
ShowImageIDAsFilename
| Tags: | @UR023-1
|  | Given | I have submitted an image for analysis
|  | When | the image is displayed
|  | Then | the ID of the image is shown
|  | And | set as the image filename

## MAGNIFYIMAGE
MagnifyImage
| Tags: | @UR, @UR004, @Analysis
As a dentist I want to zoom in on a selected area of the image So that I can see potentially affected areas in more detail
MagnifierOnImage
| Tags: | @UR004-1
|  | Given | I submit an image
|  | And | the image is displayed on screen
|  | When | I hold the left mouse button down
|  | Then | a portion of the image is displayed at a higher magnification
HideAnnotationOnMagnify
| Tags: | @UR004-2
|  | Given | I submit an image
|  | And | I activate the magnifier
|  | When | I move the mouse cursor over a caries ROI indicator
|  | Then | the indicator is hidden in the magnified view
|  | When | I move the mouse cursor over a tooth indicator
|  | Then | the indicator is hidden in the magnified view

## RADIOGRAPHQUALITYGRADING
RadiographQualityGrading
| Tags: | @UR, @UR040, @Analysis
As a dentist I want AssistDent to grade the radiograph quality So that I can be warned if AssistDent's analysis is unreliable.
ResolutionRadiograph
| Tags: | @UR040-01
|  | Given | I have submitted a radiograph for analysis
|  | When | the radiograph has been analysed
|  | Then | I see the width and height displayed within the radiograph quality report
|  | And | I see a rating label for the resolution within the radiograph quality report
|  | And | the status for the resolution is displayed within the radiograph quality report
GrayScaleRadiograph
| Tags: | @UR040-02
|  | Given | I have submitted a radiograph for analysis
|  | When | the radiograph has been analysed
|  | Then | I see the gray scale flag displayed within the radiograph quality report
|  | And | the status for the gray scale flag is displayed within the radiograph quality report
ContrastRadiograph
| Tags: | @UR040-03
|  | Given | I have submitted a radiograph for analysis
|  | When | the radiograph has been analysed
|  | Then | I see the contrast metric displayed within the radiograph quality report
|  | And | I see a rating label for the contrast metric within the radiograph quality report
|  | And | the status for the contrast metric is displayed within the radiograph quality report
BlurRadiograph
| Tags: | @UR040-04
|  | Given | I have submitted a radiograph for analysis
|  | When | the radiograph has been analysed
|  | Then | I see the blur metric displayed within the radiograph quality report
|  | And | I see a rating label for the blur metric within the radiograph quality report
|  | And | the status for the blur metric is displayed within the radiograph quality report

## REPOSITIONCARIESROI
RepositionCariesROI
| Tags: | @UR, @UR009, @Analysis
As a dentist I want to be able to reposition and change arrow indicator directions so that I can indicate alternative regions that I believe to be carious
MoveCariesROI
| Tags: | @UR009-1
|  | Given | I have submitted an image
|  | And | the image contains at least one ROI
|  | When | I click-and-drag with the mouse button on an ROI
|  | Then | the ROI is moved to the the new location
|  | And | the caries grade label is moved to the new location
|  | And | the ROI points away from the tooth centre to the correct edge
|  | And | the caries tooth location text is updated in the text report
|  | And | the mesial/distal caries text is updated in the text report
FlipCariesROIUsingButton
| Tags: | @UR009-2
|  | Given | I submit an image
|  | And | the image contains caries ROIs
|  | When | I select one or more caries ROIs
|  | And | flip its direction
|  | Then | the caries ROIs direction is flipped horizontally
|  | And | the caries tooth location text is updated in the text report
|  | And | the mesial/distal caries text is updated in the text report
FlipCariesROIUsingMouse
| Tags: | @UR009-3
|  | Given | I submit an image
|  | And | the image contains caries ROIs
|  | When | I click-drag-release with the mouse button on an ROI
|  | And | the new location is the same tooth but different edge
|  | Then | the caries ROIs arrow direction is flipped horizontally
|  | And | the ROI points away from the tooth centre to the correct edge
|  | And | the caries tooth location text is updated in the text report
|  | And | the mesial/distal caries text is updated in the text report

## SHOWPROGRESS
ShowProgress
| Tags: | @UR, @UR026, @Analysis
As a user I want confirmation that AssistDent is performing caries detection So that I can wait for the analysis to complete
ShowAnalysisProgress
| Tags: | @UR026-1
|  | When | I perform caries detection
|  | Then | the analysis status is reported at relevant stages

## TEXTREPORTEXPORT
TextReportExport
| Tags: | @UR, @UR039, @Analysis
As a dentist I want AssistDent to export the analysis results in text format So that I can store the text report alongside an exported image
ExportTextReportOnDemand
| Tags: | @UR039-01
|  | Given | I have performed caries detection
|  | When | I export an image
|  | Then | a text report file is created
|  | And | that file contains a text report corresponding to the exported image
|  | And | the text report contains a warning to review the exported image
ExportTextReportAutomatically
| Tags: | @UR039-02
|  | Given | I have performed caries detection
|  | And | I have enabled automatic export
|  | When | I close AssistDent
|  | Then | an text report file is created
|  | And | that file contains a text report corresponding to the exported image
|  | And | the text report contains a warning to review the exported image

## TOGGLEANNOTATIONVISIBILITY
ToggleAnnotationVisibility
| Tags: | @UR, @UR007, @Analysis
As a dentist I want to hide all annotation So that I can review the image grade and look for other conditions
HideImageAnnotation
| Tags: | @UR007-1
|  | Given | I submit an image
|  | And | the image contains caries ROIs
|  | When | I hide the annotation
|  | Then | the tooth and caries ROI indicators are hidden from view
RestoreImageAnnotation
| Tags: | @UR007-2
|  | Given | I submit an image
|  | And | the image contains caries ROIs
|  | And | I have hidden the annotation
|  | When | I restore the annotation
|  | Then | the tooth and caries ROI indicators are displayed