## TOC
user
AddCariesROI
DeleteCariesROI
DesktopUsageMetrics
DisplayCariesDetection
DisplayImage
DisplayTeethDetection
DragAndDropExport
DragAndDropImport
EnableDentine
ImageExport
MagnifyImage
RepositionCariesROI
ShowProgress
ToggleAnnotationVisibility

## ADDCARIESROI
AddCariesROI
| Tags: | @UR, @UR008, @Analysis
As a dentist I want to add caries regions of interest at a location I deem dubious So that I can record for future use and share the results with a patient
AddCariesROI
| Tags: | @UR008-1
|  | Given | I submit an image
|  | When | I choose a location at which to add a caries ROI to the image
|  | Then | the new caries ROI is show on the image
|  | And | the new caries ROI is labled as of unknown grade

## DELETECARIESROI
DeleteCariesROI
| Tags: | @UR, @UR006, @Analysis
As a dentist I want to delete caries regions of interest that I judge to be incorrect So that I can share the results with a patient without showing false positives
DeleteCariesROI
| Tags: | @UR006-1
|  | Given | I submit an image
|  | And | the image contains caries ROIs
|  | When | I select one or more caries ROIs
|  | And | delete them
|  | Then | the caries ROIs are removed from the image

## DESKTOPUSAGEMETRICS
DesktopUsageMetrics
| Tags: | @UR, @UR032, @DesktopOnly, @Analysis
As a company We want to record the number of times AssistDent is used and the date of last use So that we can monitor take up of the product during evaluations
RecordNumberOfUses
| Tags: | @UR032-01
|  | Given | I have registered
|  | And | used AssistDent n times
|  | When | I view the information associated with my licence that incorporates my registered name
|  | Then | that information shows AssistDent was used n times
RecordNumberOfUses
| Tags: | @UR032-02
|  | Given | I have registered
|  | And | used AssistDent today
|  | When | I view the information associated with my licence that incorporates my registered name
|  | Then | that information shows AssistDent was last used today

## DISPLAYCARIESDETECTION
DisplayCariesDetection
| Tags: | @UR, @UR003, @Analysis
As a dentist I want to view caries regions of interest on a bitewing image So that I can see where caries is ocurring
ROIIndicatorsAreDisplayed
| Tags: | @UR003-1
|  | When | I submit an image
|  | Then | the detected caries ROIs are indicated
ROIIndicatorsAreEasilyVisible
| Tags: | @UR003-2
|  | When | I submit an image
|  | Then | the detected caries ROIs are indicated
|  | And | the detected caries ROI indicators are easily visible
|  | And | the detected caries ROI indicators for different grades are distinguishable

## DISPLAYIMAGE
DisplayImage
| Tags: | @UR, @UR001, @Analysis
As a dentist I want to view one or more bitewing images So that when caries detection is performed I can view the results in place over the analyzed images
ImageIsDisplayed
| Tags: | @UR001-1
|  | When | I submit an image
|  | Then | the image is displayed on screen

## DISPLAYTEETHDETECTION
DisplayTeethDetection
| Tags: | @UR, @UR002, @Analysis
As a dentist I want to see which teeth have been detected and analyzed for caries So that I can check for caries manually on undetected teeth
TeethIndicatorsAreDisplayed
| Tags: | @UR002-1
|  | When | I submit an image
|  | Then | the detected teeth are indicated on the image
|  | And | idenfiers for the teeth are shown
EditToothQuadrant
| Tags: | @UR002-2, @DesktopOnly
|  | Given | teeth identifiers are show
|  | Then | I can edit the tooth quadrant
EditToothNumber
| Tags: | @UR002-3, @DesktopOnly
|  | Given | teeth identifiers are show
|  | Then | I can edit the tooth number

## DRAGANDDROPEXPORT
DragAndDropExport
| Tags: | @UR, @UR036, @DesktopOnly, @Analysis
As a dentist I want to drag an image, including annotation, from AssistDent into my dental imaging system So that I can keep a record of the caries detection results
ExportByDragAndDrop
| Tags: | @UR036-01
|  | Given | I have analyzed an image in AssistDent
|  | When | I drag the image out of AssistDent
|  | And | drop the image into another application that supports the drop operation
|  | Then | the image is loaded into the other application

## DRAGANDDROPIMPORT
DragAndDropImport
| Tags: | @UR, @UR035, @DesktopOnly, @Analysis
As a dentist I want to drag an image from my dental imaging system into AssistDent So that I can perform caries detection on my patients images
ImportByDragAndDrop
| Tags: | @UR035-01
|  | Given | AssistDent is running
|  | When | I drag an image from another application
|  | And | drop the image into AssistDent
|  | Then | AssistDent loads and analyzes the dropped image

## ENABLEDENTINE
EnableDentine
| Tags: | @UR, @UR037, @DesktopOnly, @Analysis
As a dentist I want to enable and disable the display and incorporation of dentine caries in my analysis So that I can optionally concentrate on just enamel-only proximal caries
EnableDentine
| Tags: | @UR037-01
|  | Given | I have enabled dentine
|  | When | I open an image which contains enamel and dentine caries
|  | Then | both enamel and dentine caries are displayed
|  | And | I can change the grade of a caries roi to D3 or D4
DisableDentine
| Tags: | @UR037-02
|  | Given | I have disabled dentine
|  | When | I open an image which contains enamel and dentine caries
|  | Then | only enamel caries are displayed
|  | And | I am unable to change the grade to D3 or D4

## IMAGEEXPORT
ImageExport
| Tags: | @UR, @UR031, @DesktopOnly, @Analysis
As a dentist I want AssistDent to export the image with annotated results So that I can share the image with the patient and/or import it into my dental imaging software
ExportOnDemand
| Tags: | @UR031-01
|  | Given | I have performed caries detection
|  | When | I export an image
|  | Then | an image file is created
|  | And | that image contains the source X-Ray and annoted caries ROIs
ExportAutomatically
| Tags: | @UR031-02
|  | Given | I have performed caries detection
|  | And | I have enabled automatic export
|  | When | I close AssistDent
|  | Then | an image file is created
|  | And | that image contains the source X-Ray and annoted caries ROIs
SnapshotImage
| Tags: | @UR031-03
|  | Given | I have performed caries detection
|  | When | I snapshot an image
|  | Then | the clipboard contains an image
|  | And | that image contains the source X-Ray and annoted caries ROIs
RunCommandAfterExport
| Tags: | @UR031-04
|  | Given | I have performed caries detection
|  | And | I have configured AssistDent to run a command after export
|  | When | I export an image
|  | Then | the export is created
|  | And | the post-export command is executed

## MAGNIFYIMAGE
MagnifyImage
| Tags: | @UR, @UR004, @Analysis
As a dentist I want to zoom in on a selected area of the image So that I can see potentially affected areas in more detail
MagnifierOnImage
| Tags: | @UR004-1
|  | Given | I submit an image
|  | And | the image is displayed on screen
|  | When | I hold the left mouse button down
|  | Then | a portion of the image is displayed at a higher magnification
HideAnnotationOnMagnify
| Tags: | @UR004-2
|  | Given | I submit an image
|  | And | I activate the magnifier
|  | When | I move the mouse cursor over a caries ROI indicator
|  | Then | the indicator is hidden in the magnified view
|  | When | I move the mouse cursor over a tooth indicator
|  | Then | the indicator is hidden in the magnified view

## REPOSITIONCARIESROI
RepositionCariesROI
| Tags: | @UR, @UR009, @Analysis
As a dentist I want to be able to reposition and change arrow indicator directions so that I can indicate alternative regions that I believe to be carious
MoveCariesROI
| Tags: | @UR009-1
|  | Given | I have submitted an image
|  | And | the image contains at least one ROI
|  | When | I click-and-drag with the mouse button on an ROI
|  | Then | the ROI is moved to the the new location
|  | And | the caries grade label is moved to the new location
FlipCariesROI
| Tags: | @UR009-3
|  | Given | I submit an image
|  | And | the image contains caries ROIs
|  | When | I select one or more caries ROIs
|  | And | flip its direction
|  | Then | the caries ROIs direction is flipped horizontally
|  | And | the caries grade label is in the correct location

## SHOWPROGRESS
ShowProgress
| Tags: | @UR, @UR026, @Analysis
As a user I want confirmation that AssistDent is performing caries detection So that I can wait for the analysis to complete
ShowAnalysisProgress
| Tags: | @UR026-1
|  | When | I perform caries detection
|  | Then | the analysis status is reported at relevant stages

## TOGGLEANNOTATIONVISIBILITY
ToggleAnnotationVisibility
| Tags: | @UR, @UR007, @Analysis
As a dentist I want to hide all annotation So that I can review the image grade and look for other conditions
HideImageAnnotation
| Tags: | @UR007-1
|  | Given | I submit an image
|  | And | the image contains caries ROIs
|  | When | I hide the annotation
|  | Then | the tooth and caries ROI indicators are hidden from view
RestoreImageAnnotation
| Tags: | @UR007-2
|  | Given | I submit an image
|  | And | the image contains caries ROIs
|  | And | I have hidden the annotation
|  | When | I restore the annotation
|  | Then | the tooth and caries ROI indicators are displayed