## TOC
functional
AddROI
DeleteROI
DisplayImage
DragAndDropExport
DragAndDropImport
EnableDentine
ImageExport
ImageMagnifer
ProgressDisplay
RepositionROI
ROIIndicator
ToggleAnnotation
ToothIndicator

## ADDROI
AddROI
| Tags: | @FR, @FR008, @UR008, @Analysis
AddCariesROICursorAppearanceChange
| Tags: | @FR008-1, @UR008-1
|  | Given | I have submitted an image
|  | When | click the Add ROI button
|  | Then | the Add ROI button changes appearance
|  | And | the Cursor changes appearance to a crosshair
AddCariesROICursorAppearanceReturn
| Tags: | @FR008-2, @UR008-1
|  | Given | I have selected the Add ROI Button
|  | When | I hit select the Add ROI button again or any other button
|  | Then | the Add ROI button changes appearance back to original
|  | And | the Cursor changes appearance back to default
AddCariesROI
| Tags: | @FR008-3, @UR008-1
|  | Given | I have selected the Add ROI Button
|  | When | I move the cursor to the image and click
|  | Then | a new ROI is display with the arrow point at the mouse click
|  | And | pointing away from the nearest tooth centre
|  | And | the grade of the ROI is unknown
|  | And | the Add ROI button changes appearance back to original
|  | And | the Cursor changes appearance back to default

## DELETEROI
DeleteROI
| Tags: | @FR, @FR006, @UR006, @Analysis
DeleteROI
| Tags: | @FR006-2, @UR006-1
|  | Given | an analyzed bitewing image which contains caries ROIs
|  | When | I select an ROI
|  | And | click the delete button corresponding to that image
|  | Then | the selected caries ROI will be deleted
DeleteROIs
| Tags: | @FR006-3, @UR006-1
|  | Given | an analyzed bitewing image which contains caries ROIs
|  | When | I select 2 or more ROIs
|  | And | click the delete button corresponding to that image
|  | Then | the selected caries ROIs will be deleted

## DISPLAYIMAGE
DisplayImage
| Tags: | @FR, @FR001, @UR001, @Analysis
The image shall be displayed on screen
DisplayImage
| Tags: | @FR001-1, @UR001-1
|  | Given | a bitewing image
|  | When | submitted for analysis
|  | Then | the bitewing image is displayed on screen

## DRAGANDDROPEXPORT
DragAndDropExport
| Tags: | @FR, @FR038, @UR036, @Analysis, @DesktopOnly
ShowDragCursor
| Tags: | @FR037-01, @UR036-01
|  | Given | an image has been analyzed in AssistDent
|  | When | I click and drag the right-mouse button
|  | Then | the cursor is updated to show I am dragging something
DraggedExportFilename
| Tags: | @FR037-02, @UR036-01
|  | Given | an image has been analyzed in AssistDent
|  | When | I click and drag the right-mouse button
|  | And | drop the image onto an app which supports drag and drop operations
|  | Then | a file appears in the app
|  | And | the filename incorporates the filename of the original image file
CanDropIntoOtherAppEnamelOnly
| Tags: | @FR037-03, @UR036-01
|  | Given | an image has been analyzed in AssistDent which contains enamel and dentine caries
|  | And | include dentine caries is unchecked
|  | When | I click and drag the right-mouse button
|  | And | drop the image onto an app which supports drag and drop operations
|  | Then | a file appears in the app
|  | And | the file contains the x-ray image
|  | And | the file contains the enamel caries annotations and NOT the dentine annotations
CanDropIntoOtherAppEnamelAndDentine
| Tags: | @FR037-04, @UR036-01
|  | Given | an image has been analyzed in AssistDent which contains enamel and dentine caries
|  | And | include dentine caries is checked
|  | When | I click and drag the right-mouse button
|  | And | drop the image onto an app which supports drag and drop operations
|  | Then | a file appears in the app
|  | And | the file contains the x-ray image
|  | And | the file contains the enamel caries annotations and the dentine annotations

## DRAGANDDROPIMPORT
DragAndDropImport
| Tags: | @FR, @FR037, @UR035, @Analysis, @DesktopOnly
AcceptDraggedImage
| Tags: | @FR037-01, @UR035-01
|  | Given | AssistDent is started
|  | When | I drag and image from another application
|  | And | drop the image into AssistDent
|  | Then | the image is loaded and analyzed
HandleNonImage
| Tags: | @FR037-02, @UR035-01
|  | Given | AssistDent is started
|  | When | I drag a non-image file from another application
|  | And | drop the file into AssistDent
|  | Then | nothing happens
IgnoreMultipleImages
| Tags: | @FR037-03, @UR035-01
|  | Given | AssistDent is started
|  | When | I drag 2 or more images from another application
|  | And | drop the files into AssistDent
|  | Then | nothing happens

## ENABLEDENTINE
EnableDentine
| Tags: | @FR, @FR039, @UR037, @Analysis, @DesktopOnly
ChangeDentineDisplay
| Tags: | @FR039-01, @UR036-01
|  | Given | I open the settings window
|  | And | change the option to include dentine caries
|  | When | I restart AssistDent
|  | And | open the settings window
|  | Then | my preference has been remembered
ShowDentineCaries
| Tags: | @FR039-02, @UR036-01
|  | Given | I have enabled the display of caries using the settings window
|  | When | I open an image containing both enamel and dentine caries
|  | Then | the dentine caries are indicated
|  | And | the enamel caries are indicated
HideDentineCaries
| Tags: | @FR039-03, @UR036-02
|  | Given | I have disabled the display of caries using the settings window
|  | When | I open an image containing both enamel and dentine caries
|  | Then | the dentine caries are not indicated
|  | And | the enamel caries are indicated
AllowAssignmentOfDentineGrades
| Tags: | @FR039-04, @UR036-01
|  | Given | I have enabled the display of dentine caries using the settings window
|  | And | I open an image containing both enamel and dentine caries
|  | And | dentine caries are indicated
|  | When | I left-click on a dentine caries roi
|  | Then | the grades menu is displayed
|  | And | the grade can be set to D
DenyAssignmentOfDentineGrades
| Tags: | @FR039-04, @UR036-02
|  | Given | I have disabled the display of dentine caries using the settings window
|  | And | I open an image containing enamel caries
|  | When | I left-click on a caries roi
|  | Then | the grades menu is displayed
|  | And | the grade cannot be set to D

## IMAGEEXPORT
ImageExport
| Tags: | @FR, @FR034, @UR031, @Analysis, @DesktopOnly
ExportImageButton
| Tags: | @FR034-01, @UR031-01
|  | Given | I have performed caries detection
|  | When | I click the Export button
|  | Then | an image file is created
|  | And | that image contains the source X-Ray and annotated caries ROIs
ExportROIsEnamelOnly
| Tags: | @FR034-02, @UR031-01
|  | Given | I have performed caries detection on an image containing dentine caries
|  | And | include dentine caries is unchecked
|  | When | I click the Export button
|  | Then | an image file is created
|  | And | that image contains the source X-Ray and annotated enamel caries ROIs
|  | And | dentine caries ROI are NOT displayed
ExportROIsEnamelAndDentine
| Tags: | @FR034-03, @UR031-01
|  | Given | I have performed caries detection on an image containing dentine caries
|  | And | include dentine caries is checked
|  | When | I click the Export button
|  | Then | an image file is created
|  | And | that image contains the source X-Ray and annotated enamel caries ROIs
|  | And | dentine caries ROI are displayed
ExportToNewLocation
| Tags: | @FR034-04, @UR031-01
|  | Given | I have performed caries detection
|  | And | I configure the export to be created in a new folder
|  | When | I click the Export button
|  | Then | an image file is created in the new location
|  | And | that image contains the source X-Ray and annoted caries ROIs
ExportHasUniqueName
| Tags: | @FR034-05, @UR031-01
|  | Given | I have performed caries detection
|  | When | I click the Export button
|  | Then | an image file is created
|  | And | that image file has a unique name based on the date and time
|  | And | that name of the original input image is included in the name
ExportCreatedOnExit
| Tags: | @FR034-06, @UR031-02
|  | Given | I have performed caries detection
|  | And | I have configured AssistDent to export the image on exit
|  | When | I close AssistDent
|  | Then | an image file is created
|  | And | that image contains the source X-Ray and annoted caries ROIs
SnapshotImage
| Tags: | @FR034-07, @UR031-03
|  | Given | I have performed caries detection
|  | When | I click the Snapshot button
|  | Then | an image file is copied to the clipboard
|  | And | that image contains the source X-Ray and annoted caries ROIs
RunCommandAfterExport
| Tags: | @FR034-08, @UR031-04
|  | Given | I have performed caries detection
|  | And | I have configured AssisDent to run a command after export
|  | When | I click the Export button
|  | Then | the export is created
|  | And | the configured command is executed
ExportToVistaSoft
| Tags: | @FR034-09, @UR031-04
|  | Given | I have performed caries detection
|  | And | I have configured AssisDent to send to VistaSoft after export
|  | When | I click the Export button
|  | Then | the export is created
|  | And | the image appears alongside the parent image in VistaSoft
ExportAs
| Tags: | @FR034-10, @UR031-01
|  | Given | I have performed caries detection
|  | When | I click the ExportAs option
|  | Then | the file save dialog opens
|  | And | the export image file is created
|  | And | the export image file is in the specified folder with the specified filename
ExportAsCancel
| Tags: | @FR034-11, @UR031-01
|  | Given | I have performed caries detection
|  | When | I click the ExportAs option
|  | And | I choose the cancel button
|  | Then | no export image file is created

## IMAGEMAGNIFER
ImageMagnifer
| Tags: | @FR, @FR004, @UR004, @Analysis
With the mouse button held down, the section of the image directly under the mouse pointer will be zoomed to 3x magnification
ShowMagnifer
| Tags: | @FR004-1, @UR004-1
|  | Given | a bitewing image displayed on screen
|  | When | I hover the mouse pointer over a location in the image
|  | And | I hold the mouse button down
|  | Then | the magnifier should be displayed
|  | And | the image location should be displayed at 3x magnification
MoveMagnifer
| Tags: | @FR004-2, @UR004-1
|  | Given | a bitewing image displayed on screen
|  | And | the magnifier is currently displayed
|  | When | I move the mouse pointer to another location in the image
|  | Then | the image location should be updated
HideAnnotationInMagnifer
| Tags: | @FR004-3, @UR004-2
|  | Given | a bitewing image displayed on screen
|  | And | the magnifier is currently displayed
|  | And | the magnifer is positioned over an annotation
|  | Then | no annotation should be displayed in the magnification

## PROGRESSDISPLAY
ProgressDisplay
| Tags: | @FR, @FR028, @UR026, @Analysis, @DesktopOnly
ShowAnalysisFeedback
| Tags: | @FR028-1, @UR026-1
|  | When | I submit an image for analysis
|  | Then | the app displays status updates
|  | And | the updates are clearly visible

## REPOSITIONROI
RepositionROI
| Tags: | @FR, @FR009, @UR009, @Analysis
MoveAndDragROI
| Tags: | @FR009-1, @UR009-1
|  | Given | an analyzed bitewing image which contains caries ROIs
|  | When | I click and drag an ROI
|  | Then | the ROI and caries grade label will move with the mouse
|  | And | the ROI selection status will remain unchanged
MoveROI
| Tags: | @FR009-2, @UR009-1
|  | Given | an analyzed bitewing image which contains caries ROIs
|  | When | I click and drag an ROI
|  | And | move the mouse
|  | And | release the mouse
|  | Then | the ROI and grade label will stay at its new location
|  | And | the ROI selection status will remain unchanged
FlipROI
| Tags: | @FR009-3, @UR009-2
|  | Given | an analyzed bitewing image which contains caries ROIs
|  | When | I select an ROI
|  | And | click the flip button corresponding to that image
|  | Then | the selected caries ROI will be flipped
|  | And | the ROI will remain selected
FlipROIs
| Tags: | @FR009-4, @UR009-2
|  | Given | an analyzed bitewing image which contains caries ROIs
|  | When | I select 2 or more ROIs
|  | And | click the flip button corresponding to that image
|  | Then | the selected caries ROIs will be flipped
|  | And | the ROIs will remain selected

## ROIINDICATOR
ROIIndicator
| Tags: | @FR, @FR003, @UR003, @Analysis
The location of caries regions of interest shall be displayed in position on the image
ROIIndicators
| Tags: | @FR003-1, @UR003-1
|  | Given | a bitewing image known to contain caries
|  | When | caries detection is complete
|  | Then | caries ROIs are indicated over the image
ROIArrows
| Tags: | @FR003-2, @UR003-1
|  | Given | a bitewing image known to contain caries
|  | When | caries detection is complete
|  | Then | caries ROIs are indicated by arrows
|  | And | the arrows do not obscure the tooth edge
ROIEasilySeen
| Tags: | @FR003-3, @UR003-1
|  | Given | a bitewing image known to contain caries
|  | When | caries detection is complete
|  | Then | caries ROI indicators are a consistent size
|  | And | caries ROI indicators are brightly coloured
ROIGrades
| Tags: | @FR003-5, @UR003-2
|  | Given | a bitewing image known to contain caries of different grades
|  | When | caries detection is complete
|  | Then | caries ROI indicators for different grades are distinguishable
EditCariesGrade
| Tags: | @FR003-6, @UR003-2, @DesktopOnly
|  | Given | caries ROI are displayed on the image
|  | When | I click on the ROI label
|  | Then | a menu allowing me to change the grade appearns
|  | And | I can choose a new grade
|  | And | the grade is updated with the new choice

## TOGGLEANNOTATION
ToggleAnnotation
| Tags: | @FR, @FR007, @UR007, @Analysis
HideImageAnnotation
| Tags: | @FR007-2, @UR007-1
|  | Given | an analyzed bitewing image which contains caries ROIs
|  | When | I click the corresponding hide annotation button
|  | Then | the tooth indicators are hidden
|  | And | the caries ROI indicators are hidden
|  | And | the hide annotation button is replaced with a show annotation button
RestoreImageAnnotation
| Tags: | @FR007-3, @UR007-2
|  | Given | an analyzed bitewing image which contains caries ROIs
|  | And | the annotation is currently hidden
|  | When | I click the corresponding show annotation button
|  | Then | the tooth indicators are displayed
|  | And | the caries ROI indicators are displayed
|  | And | the show annotation button is replaced with a hide annotation button

## TOOTHINDICATOR
ToothIndicator
| Tags: | @FR, @FR002, @UR002, @Analysis
Detected teeth shall be displayed in position over the image
ToothIndicators
| Tags: | @FR002-1, @UR002-1
|  | Given | a bitewing image known to contain adult teeth
|  | When | caries detection is complete
|  | Then | the detected teeth are displayed on the image
|  | And | identifiers for the teeth are shown
|  | And | the identifiers provide unique identification for each tooth
EditToothQuadrant
| Tags: | @FR002-2, @UR002-2, @DesktopOnly
|  | Given | identifier for the teeth are displayed on the image
|  | When | I click on the tooth quadrant
|  | Then | a menu allowing me to change the quadrant appearns
|  | And | I can choose a new quadrant
|  | And | the label is updated with the new quadrant value
EditToothNumber
| Tags: | @FR002-3, @UR002-3, @DesktopOnly
|  | Given | identifier for the teeth are displayed on the image
|  | When | I click on the tooth number
|  | Then | a menu allowing me to change the number appearns
|  | And | I can choose a new number
|  | And | the label is updated with the new number value